## Summary
Summary
Subject:
Laboratory:
Result | Date | Laboratory | Material |  | Method X |  |  |  | Method Y | [6 empty] | Passing / Bablok |  |  | Disp. of Residuals |  |  | Kendall's
Sheet |  |  |  | Short Name | Long Name | Unit | Instrument | Short Name | Long Name | Unit | Instrument | n | Min X | Max X | B/P b | sign. (b) | B/P a | sign. (a) | md(68) | md(95) | r | Tau | Linearity
[13 empty] | 0 | 0 | 0 |  | 0 |  | 0 | 0 | 0 | 0

## Input Sheet
|  |  |  | Worksheet for Method Comparison | [40 empty] | 20
Subject:
Lab.:
Serial No. | Identifier | F1 | F2 | Vector 1 | Vector 2 | Vector 3 | Vector 4 | Vector 5 | Vector 6 | Vector 7 | Vector 8 | Vector 9 | Vector 10 | Vector 11 | Vector 12 | Vector 13 | Vector 14 | Vector 15 | Vector 16 | Vector 17 | Vector 18 | Vector 19 | Vector 20 | Vector 21 | Vector 22 | Vector 23 | Vector 24 | Vector 25 | Vector 26 | Vector 27 | Vector 28 | Vector 29 | Vector 30 | Vector 31 | Vector 32
| Short Name |  |  | CEA | CEA
Long Name
| [Unit] |  |  | ng/ml | ng/ml
| Instrument |  |  | E 170 | Centaur
lfd_no | ident_no | filter1 | filter2 | X_Vektor | Y_Vektor | [31 empty] | sel_X_vector | sel_Y_vector | sel_seq | sel_ident | sel_f1 | sel_f2
1 | 37247241 |  |  | 1.3 | 1.56 | [37 empty] | SumData
2 | 37248667 |  |  | 1 | 0.93 | [37 empty] | 3.1 | ver_Info
3 | 37226100 |  |  | 93.6 | 65.02 | [37 empty] | 2012-04-26 11:01 | L_DateTime
4 | 37248732 |  |  | 4.7 | 3.66 | [36 empty] | Kriterienfeld | CEA | L_SName_X
5 | 37247206 |  |  | 1.6 | 1.37 | [36 empty] | #NAME? |  | L_LName_X
6 | 32127950 |  |  | 1.8 | 1.4 | [37 empty] | ng/ml | L_Unit_X
7 | 32127940 |  |  | 3 | 2.38 | [37 empty] | E 170 | L_Instr_X
8 | 32127943 |  |  | 0.9 | 1.04 | [31 empty] | Zielbereich | [5 empty] | CEA | L_SName_Y
9 | 37241267 |  |  | 3.2 | 2.61 | [38 empty] | L_LName_Y
10 | 37249618 |  |  | 1.9 | 1.94 | [31 empty] | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 | ng/ml | L_Unit_Y
11 | 37249443 |  |  | 5 | 3.51 | [30 empty] | sel_first_row | 1.3 | 1.56 | 1 | 37247241 |  |  | Centaur | L_Instr_Y
12 | 37248600 |  |  | 241.3 | 269.7 | [31 empty] | 1 | 0.93 | 2 | 37248667 |  |  |  | L_N
13 | 37248597 |  |  | 580.1 | [32 empty] | 93.6 | 65.02 | 3 | 37226100 |  |  |  | L_Min_X
14 | 37250325 |  |  | 4.8 | 3.26 | [31 empty] | 4.7 | 3.66 | 4 | 37248732 |  |  |  | L_Max_X
15 | 37250153 |  |  | 4.4 | 3 | [31 empty] | 1.6 | 1.37 | 5 | 37247206 |  |  |  | L_PBa
16 | 37242476 |  |  | 13.7 | 9.77 | [31 empty] | 1.8 | 1.4 | 6 | 32127950 |  |  |  | L_PBas
17 | 37250366 |  |  | 427.2 | [32 empty] | 3 | 2.38 | 7 | 32127940 |  |  |  | L_PBb
18 | 32161366 |  |  | 2.9 | 2.85 | [31 empty] | 0.9 | 1.04 | 8 | 32127943 |  |  |  | L_PBbs
19 | 37250080 |  |  | 0.8 | 1.03 | [31 empty] | 3.2 | 2.61 | 9 | 37241267 |  |  |  | L_md65
20 | 38159530 |  |  | 2.1 | 1.79 | [31 empty] | 1.9 | 1.94 | 10 | 37249618 |  |  |  | L_md95
21 | 32023967 |  |  | 0.9 | 1.11 | [31 empty] | 5 | 3.51 | 11 | 37249443 |  |  |  | L_r
22 | 37221545 |  |  | 112 | 73.13 | [31 empty] | 241.3 | 269.7 | 12 | 37248600 |  |  |  | L_Tau
23 | 38160005 |  |  | 0.5 | 0.72 | [31 empty] | 4.8 | 3.26 | 14 | 37250325 |  |  |  | L_linear
24 | 37218607 |  |  | 11.4 | 6.15 | [31 empty] | 4.4 | 3 | 15 | 37250153 |  |  | 1 | dlg_x_vector
25 | 37251080 |  |  | 256.9 | 245.22 | [31 empty] | 13.7 | 9.77 | 16 | 37242476 |  |  | 2 | dlg_y_vector
26 | 37250351 |  |  | 2.8 | 1.87 | [31 empty] | 2.9 | 2.85 | 18 | 32161366 |  |  | 1 | dlg_Task_sel
27 | 37249620 |  |  | 25.4 | 13.18 | [31 empty] | 0.8 | 1.03 | 19 | 37250080 |  |  | 2 | dlg_plot_name
28 | 37251452 |  |  | 2.4 | 1.53 | [31 empty] | 2.1 | 1.79 | 20 | 38159530 |  |  | TRUE | dlg_plot_do_unit
29 | 37131730 |  |  | 1.5 | 1.37 | [31 empty] | 0.9 | 1.11 | 21 | 32023967 |  |  | TRUE | dlg_plot_do_instr
30 | 37249575 |  |  | 118 | 25.67 | [31 empty] | 112 | 73.13 | 22 | 37221545 |  |  |  | Dlg_Min_ID_No
31 | 37249589 |  |  | 85.5 | 60.75 | [31 empty] | 0.5 | 0.72 | 23 | 38160005 |  |  |  | Dlg_Max_ID_No
32 | 37252031 |  |  | 2 | 1.77 | [31 empty] | 11.4 | 6.15 | 24 | 37218607 |  |  |  | Dlg_Min_Identification
33 | 37251102 |  |  | 87 | 75.26 | [31 empty] | 256.9 | 245.22 | 25 | 37251080 |  |  |  | Dlg_Max_Identification
34 | 37221655 |  |  | 3.6 | 2.55 | [31 empty] | 2.8 | 1.87 | 26 | 37250351 |  |  |  | Dlg_Min_X_Value
35 | 37229749 |  |  | 52.7 | 29.74 | [31 empty] | 25.4 | 13.18 | 27 | 37249620 |  |  |  | Dlg_Max_X_Value
36 | 32147385 |  |  | 105.5 | 87.82 | [31 empty] | 2.4 | 1.53 | 28 | 37251452 |  |  |  | Dlg_Min_Y_Value
37 | 32127984 |  |  | 0.6 | 0.63 | [31 empty] | 1.5 | 1.37 | 29 | 37131730 |  |  |  | Dlg_Max_Y_Value
38 | 38160622 |  |  | 1.7 | 1.36 | [31 empty] | 118 | 25.67 | 30 | 37249575 |  |  |  | Dlg_Filter_1_Selection
39 | 32127910 |  |  | 2.7 | 2.06 | [31 empty] | 85.5 | 60.75 | 31 | 37249589 |  |  |  | Dlg_Filter_2_Selection
40 | 37252285 |  |  | 23 | 15.85 | [31 empty] | 2 | 1.77 | 32 | 37252031 |  |  | 1 | Dlg_ID_No_IncExc
41 | 37251541 |  |  | 6.3 | 3.48 | [31 empty] | 87 | 75.26 | 33 | 37251102 |  |  | 1 | Dlg_Identification_IncExc
42 | 37254968 |  |  | 609.6 | [32 empty] | 3.6 | 2.55 | 34 | 37221655 |  |  | 1 | Dlg_X_Value_IncExc
43 | 37253194 |  |  | 69.4 | 11.03 | [31 empty] | 52.7 | 29.74 | 35 | 37229749 |  |  | 1 | Dlg_Y_Value_IncExc
44 | 37251546 |  |  | 121.1 | 71.79 | [31 empty] | 105.5 | 87.82 | 36 | 32147385 |  |  | 1 | Dlg_Filter1_IncExc
45 | 37257769 |  |  | 9.1 | 5.2 | [31 empty] | 0.6 | 0.63 | 37 | 32127984 |  |  | 1 | Dlg_Filter2_IncExc
46 | 32143525 |  |  | 5.1 | 4.27 | [31 empty] | 1.7 | 1.36 | 38 | 38160622 |  |  | FALSE | Dlg_ShowCal
47 | 37257292 |  |  | 286 | [32 empty] | 2.7 | 2.06 | 39 | 32127910 |  |  | FALSE | Dlg_ShowQC
48 | 37257518 |  |  | 28.9 | 35.42 | [31 empty] | 23 | 15.85 | 40 | 37252285 |  |  | 0 | dlg_scale_min
49 | 37256273 |  |  | 3.8 | [32 empty] | 6.3 | 3.48 | 41 | 37251541 |  |  | 270 | dlg_scale_max
50 | 37235534 |  |  | 4.19 | 2.79 | [31 empty] | 69.4 | 11.03 | 43 | 37253194 |  |  | 54 | dlg_scale_int
51 | 38163096 |  |  | 1.2 | 1.22 | [31 empty] | 121.1 | 71.79 | 44 | 37251546 |  |  | TRUE | dlg_show_grid
52 | 37253621 |  |  | 3.1 | 2.06 | [31 empty] | 9.1 | 5.2 | 45 | 37257769
53 | 32127967 |  |  | 4.1 | 3.46 | [31 empty] | 5.1 | 4.27 | 46 | 32143525
54 | 37232405 |  |  | 34.8 | 7.95 | [31 empty] | 28.9 | 35.42 | 48 | 37257518
55 | 37258030 |  |  | 3.7 | 2.92 | [31 empty] | 4.19 | 2.79 | 50 | 37235534
56 | 37251561 |  |  | 40.4 | 31.25 | [31 empty] | 1.2 | 1.22 | 51 | 38163096
57 | 37236755 |  |  | 1.4 | 1.44 | [31 empty] | 3.1 | 2.06 | 52 | 37253621
58 | 32127959 |  |  | 1.2 | 1.07 | [31 empty] | 4.1 | 3.46 | 53 | 32127967
59 | 32023879 |  |  | 2.6 | 2.24 | [31 empty] | 34.8 | 7.95 | 54 | 37232405
60 | 37264264 |  |  | 271.9 | [32 empty] | 3.7 | 2.92 | 55 | 37258030
61 | 37085791 |  |  | 0.7 | 0.95 | [31 empty] | 40.4 | 31.25 | 56 | 37251561
62 | 37257171 |  |  | 294.6 | [32 empty] | 1.4 | 1.44 | 57 | 37236755
63 | 37264010 |  |  | 159.5 | [32 empty] | 1.2 | 1.07 | 58 | 32127959
64 | 32160308 |  |  | 3.4 | 2.72 | [31 empty] | 2.6 | 2.24 | 59 | 32023879
65 | 32023876 |  |  | 1.3 | 1.12 | [31 empty] | 0.7 | 0.95 | 61 | 37085791
66 | 37265382 |  |  | 2.5 | 2.28 | [31 empty] | 3.4 | 2.72 | 64 | 32160308
67 | 37242762 |  |  | 96.1 | 59.61 | [31 empty] | 1.3 | 1.12 | 65 | 32023876
68 | 37258748 |  |  | 4 | 2.38 | [31 empty] | 2.5 | 2.28 | 66 | 37265382
69 | 38166192 |  |  | 0.8 | 0.97 | [31 empty] | 96.1 | 59.61 | 67 | 37242762
70 | 37259938 |  |  | 82.5 | 59.92 | [31 empty] | 4 | 2.38 | 68 | 37258748
71 | 37259923 |  |  | 219.6 | 64.51 | [31 empty] | 0.8 | 0.97 | 69 | 38166192
72 | 32143588 |  |  | 0.5 | 0.9 | [31 empty] | 82.5 | 59.92 | 70 | 37259938
73 | 38167049 |  |  | 2.2 | 1.82 | [31 empty] | 219.6 | 64.51 | 71 | 37259923
74 | 37267530 |  |  | 7.6 | 6.01 | [31 empty] | 0.5 | 0.9 | 72 | 32143588
75 | 37265007 |  |  | 92.3 | 80.75 | [31 empty] | 2.2 | 1.82 | 73 | 38167049
76 | 32137467 |  |  | 8.5 | 5.57 | [31 empty] | 7.6 | 6.01 | 74 | 37267530
77 | 38167355 |  |  | 133.3 | 63.82 | [31 empty] | 92.3 | 80.75 | 75 | 37265007
78 | 37268394 |  |  | 96.4 | 76.43 | [31 empty] | 8.5 | 5.57 | 76 | 32137467
79 | 37272594 |  |  | 6.9 | 4.27 | [31 empty] | 133.3 | 63.82 | 77 | 38167355
80 | 37266343 |  |  | 475.4 | [32 empty] | 96.4 | 76.43 | 78 | 37268394
81 | 37257202 |  |  | 201.7 | 90.25 | [31 empty] | 6.9 | 4.27 | 79 | 37272594
82 | 38169785 |  |  | 5.8 | 3.85 | [31 empty] | 201.7 | 90.25 | 81 | 37257202
83 | 37273311 |  |  | 221.3 | 67.78 | [31 empty] | 5.8 | 3.85 | 82 | 38169785
84 | 32147424 |  |  | 3.3 | 3.18 | [31 empty] | 221.3 | 67.78 | 83 | 37273311
85 | 37274821 |  |  | 1.4 | 1.46 | [31 empty] | 3.3 | 3.18 | 84 | 32147424
86 | 37272645 |  |  | 77.3 | 58.34 | [31 empty] | 1.4 | 1.46 | 85 | 37274821
87 | 37273636 |  |  | 2.3 | 1.89 | [31 empty] | 77.3 | 58.34 | 86 | 37272645
88 | 38170230 |  |  | 15.8 | 6.9 | [31 empty] | 2.3 | 1.89 | 87 | 37273636
89 | 37277819 |  |  | 117 | 98.22 | [31 empty] | 15.8 | 6.9 | 88 | 38170230
90 | 37286656 |  |  | 67.3 | 14.24 | [31 empty] | 117 | 98.22 | 89 | 37277819
91 | 38176192 |  |  | 3.5 | 2.9 | [31 empty] | 67.3 | 14.24 | 90 | 37286656
92 | 37286613 |  |  | 541.9 | [32 empty] | 3.5 | 2.9 | 91 | 38176192
93 | 37285502 |  |  | 58.7 | 28.04 | [31 empty] | 58.7 | 28.04 | 93 | 37285502
94 |  |  |  | 58.2 | 37.81 | [31 empty] | 58.2 | 37.81 | 94
95 |  |  |  | 43.3 | 24.55 | [31 empty] | 43.3 | 24.55 | 95
96 |  |  |  | 65.9 | 35.6 | [31 empty] | 65.9 | 35.6 | 96
97 |  |  |  | 60.9 | 36.01 | [31 empty] | 60.9 | 36.01 | 97
98 |  |  |  | 98 | 56.25 | [31 empty] | 98 | 56.25 | 98
99 |  |  |  | 191.1 | [32 empty] | 45.3 | 31.39 | 100
100 |  |  |  | 45.3 | 31.39 | [31 empty] | 230.4 | 65.11 | 104
101 |  |  |  | 199.3
102 |  |  |  | 304.4
103 |  |  |  | 601.8
104 |  |  |  | 230.4 | 65.11
105
106
107
108
109
110
111
112
113
114
115
116
117
118
119
120
121
122
123
124
125
126
127
128
129
130
131
132
133
134
135
136
137
138
139
140
141
142
143
144
145
146
147
148
149
150
151
152
153
154
155
156
157
158
159
160
161
162
163
164
165
166
167
168
169
170
171
172
173
174
175
176
177
178
179
180
181
182
183
184
185
186
187
188
189
190
191
192
193
194
195
196
197
198
199
200
201
202
203
204
205
206
207
208
209
210
211
212
213
214
215
216
217
218
219
220
221
222
223
224
225
226
227
228
229
230
231
232
233
234
235
236
237
238
239
240
241
242
243
244
245
246
247
248
249
250
251
252
253
254
255
256
257
258
259
260
261
262
263
264
265
266
267
268
269
270
271
272
273
274
275
276
277
278
279
280
281
282
283
284
285
286
287
288
289
290
291
292
293
294
295
296
297
298
299
300
301
302
303
304
305
306
307
308
309
310
311
312
313
314
315
316
317
318
319
320
321
322
323
324
325
326
327
328
329
330
331
332
333
334
335
336
337
338
339
340
341
342
343
344
345
346
347
348
349
350
351
352
353
354
355
356
357
358
359
360
361
362
363
364
365
366
367
368
369
370
371
372
373
374
375
376
377
378
379
380
381
382
383
384
385
386
387
388
389
390
391
392
393
394
395
396
397
398
399
400
401
402
403
404
405
406
407
408
409
410
411
412
413
414
415
416
417
418
419
420
421
422
423
424
425
426
427
428
429
430
431
432
433
434
435
436
437
438
439
440
441
442
443
444
445
446
447
448
449
450
451
452
453
454
455
456
457
458
459
460
461
462
463
464
465
466
467
468
469
470
471
472
473
474
475
476
477
478
479
480
481
482
483
484
485
486
487
488
489
490
491
492
493
494
495
496
497
498
499
500
501
502
503
504
505
506
507
508
509
510
511
512
513
514
515
516
517
518
519
520
521
522
523
524
525
526
527
528
529
530
531
532
533
534
535
536
537
538
539
540
541
542
543
544
545
546
547
548
549
550
551
552
553
554
555
556
557
558
559
560
561
562
563
564
565
566
567
568
569
570
571
572
573
574
575
576
577
578
579
580
581
582
583
584
585
586
587
588
589
590
591
592
593
594
595
596
597
598
599
600
601

## Werte bis 10
Method X: CEA ng/ml |  |  | Instrument: | E 170
Method Y: CEA ng/ml |  |  | Instrument: | Centaur
52
Descriptive Statistics
| X | Y | Y-X | (Y - X)% of X
Median | 2.55 | 2 | -0.39 | -16.7087
Mean | 2.9767 | 2.3023 | -0.67442 | -11.3726
Minimum | 0.5 | 0.63 | -3.9 | -44.7619
Maximum | 9.1 | 6.01 | 0.4 | 80
68% Median Distance | 1.7 | 1.04 | 0.63 | 18.5946
Standard Deviation | 2.0764 | 1.2656 | 0.91379 | 24.1682
Differences
Medians |  |  |  | -21.5686
Means |  |  |  | -22.6565
Regression and Correlation Analysis
Coefficients of Correlation: | 0.96621 | 0.86727 |  |  |  |  | 2.55 | 2 | -0.39 | -16.7087 |  | 3.1
|  | slope b | intercept a | lower limit | upper limit |  | 2.9767 | 2.3023 | -0.67442 | -11.3726 |  | 2012-04-25 12:25
Structural Relationship Model: | [6 empty] | 0.5 | 0.63 | -3.9 | -44.7619 |  | CEA
Passing/Bablok (P/B) |  | 0.61325 | 0.47947 |  |  |  | 9.1 | 6.01 | 0.4 | 80
95% Confidence Region for b (P/B) |  |  |  | 0.56364 | 0.6775 |  | 1.7 | 1.04 | 0.63 | 18.5946 |  | ng/ml
95% Confidence Region for a (P/B) |  |  |  | 0.3605 | 0.57955 |  | 2.0764 | 1.2656 | 0.91379 | 24.1682 |  | E 170
Std. Principal Component (SPC) |  | 0.60951 | 0.48797 |  |  |  | -21.5686 | -22.6565 | 0 | 0 |  | CEA
Linear Model: | [6 empty] | 0.96621 | 0.86727 | 0.9758 | 0
Least Squares Regression |  | 0.58891 | 0.54927 |  |  |  | 0.61325 | 1 | 0.56364 | 0.6775 |  | ng/ml
Theil Regression |  | 0.60335 | 0.49462 |  |  |  | 0.47947 | 1 | 0.3605 | 0.57955 |  | Centaur
Dispersion of Residuals: | [6 empty] | 0.60951 | 1 | 0.48797 | 0 |  | 52
Passing/Bablok Regression | md(68) = | 0.22436 | md(95) = | 0.65505 |  |  | 0.58891 | 0 | 0.54927 | 0
Std. Principal Component |  |  | SE = | 0.3192 |  |  | 1.5852 | 0 | -0.67299 | 0
[7 empty] | 0.60335 | 0 | 0.49462 | 0 |  | 0.61325
[7 empty] | 0.22436 | 0.65505 | 0.3192 | 0 |  | Yes
[7 empty] | 0 | 0 | 1 | 52 |  | 0.47947
Yes
0.22436
0.65505
0.96621
0.86727
1
2
1
2
TRUE
TRUE
0
10
0
10
1
1
1
1
1
1
FALSE
FALSE
0
100
20
TRUE
Statistics | Method X: CEA ng/ml | Method Y: CEA ng/ml
N | 52 | 52
Mean | 2.9767 | 2.3023
Median | 2.55 | 2
Minimum | 0.5 | 0.63
Maximum | 9.1 | 6.01
Range | 8.6 | 5.38
Difference Plot
(Normalized Differences)
| Sample Values | [18 empty] | 5 | 10 | [5 empty] | 5 | 20
Serial Number | X | Y | Y - X | (Y - X)% of X | Normalized Difference (%) | [14 empty] | 10 | 15 | [5 empty] | 10 | 25
[6 empty] | Diff Data Preset |  | Diff Ref Preset
scale_min | 0 | 0 | 0 |  |  | 0.4 | 0.05 | 0 | -0.05 |  |  |  | SelectData | [6 empty] | SelectCalibs | [6 empty] | SelectQCs | [6 empty] | F1_select |  | F2_select |  | grp_select
scale_max | 10 |  | 10 |  |  | 0.6 | -0.05 | 57.1429 | -0.05 |  |  |  | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? |  | #NAME?
| 0 |  |  | 0.47947 |  |  |  | -57.6687 |  |  |  | data_count | 52 | [5 empty] | calib_count | 0 | [5 empty] | QC_count | 0 | [5 empty] | F1_count |  | F2_count |  | selFilt_count |  |  |  |  | selCrit |  |  | H | I | J | K | L | M
| 10 |  |  | 6.6119 | [8 empty] | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | filter1 |  | filter2 |  | Selected Filters | F1 | Grp Start | Legend |  | X_Vektor | Y_Vektor | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2
1 | 1.3 | 1.56 | 0.26 | 20 | 18.1818 | 0 | 0.61 | 36.0656 | 1 | 1.43 | 18.1818 |  | 1.3 | 1.56 | 1 | 37247241
2 | 1 | 0.93 | -0.07 | -7 | -7.2539 | 0 | 0.615 | 4.878 | 2 | 0.965 | -7.2539 |  | 1 | 0.93 | 2 | 37248667
4 | 4.7 | 3.66 | -1.04 | -22.1277 | -24.8804 | 0 | 0.7 | 57.1429 | 3 | 4.18 | -24.8804 |  | 4.7 | 3.66 | 4 | 37248732
5 | 1.6 | 1.37 | -0.23 | -14.375 | -15.4882 | 0 | 0.825 | 30.303 | 4 | 1.485 | -15.4882 |  | 1.6 | 1.37 | 5 | 37247206
6 | 1.8 | 1.4 | -0.4 | -22.2222 | -25 | 0 | 0.885 | 19.209 | 5 | 1.6 | -25 |  | 1.8 | 1.4 | 6 | 32127950
7 | 3 | 2.38 | -0.62 | -20.6667 | -23.0483 | 0 | 0.915 | 25.1366 | 6 | 2.69 | -23.0483 |  | 3 | 2.38 | 7 | 32127940
8 | 0.9 | 1.04 | 0.14 | 15.5556 | 14.433 | 0 | 0.965 | -7.2539 | 7 | 0.97 | 14.433 |  | 0.9 | 1.04 | 8 | 32127943
9 | 3.2 | 2.61 | -0.59 | -18.4375 | -20.3098 | 0 | 0.97 | 14.433 | 8 | 2.905 | -20.3098 |  | 3.2 | 2.61 | 9 | 37241267
10 | 1.9 | 1.94 | 0.04 | 2.1053 | 2.0833 | 0 | 1.005 | 20.8955 | 9 | 1.92 | 2.0833 |  | 1.9 | 1.94 | 10 | 37249618
11 | 5 | 3.51 | -1.49 | -29.8 | -35.0176 | 0 | 1.135 | -11.4537 | 10 | 4.255 | -35.0176 |  | 5 | 3.51 | 11 | 37249443
14 | 4.8 | 3.26 | -1.54 | -32.0833 | -38.2134 | 0 | 1.21 | 1.6529 | 11 | 4.03 | -38.2134 |  | 4.8 | 3.26 | 14 | 37250325
15 | 4.4 | 3 | -1.4 | -31.8182 | -37.8378 | 0 | 1.21 | -14.876 | 12 | 3.7 | -37.8378 |  | 4.4 | 3 | 15 | 37250153
18 | 2.9 | 2.85 | -0.05 | -1.7241 | -1.7391 | 0 | 1.42 | 2.8169 | 13 | 2.875 | -1.7391 |  | 2.9 | 2.85 | 18 | 32161366
19 | 0.8 | 1.03 | 0.23 | 28.75 | 25.1366 | 0 | 1.43 | 18.1818 | 14 | 0.915 | 25.1366 |  | 0.8 | 1.03 | 19 | 37250080
20 | 2.1 | 1.79 | -0.31 | -14.7619 | -15.9383 | 0 | 1.43 | 4.1958 | 15 | 1.945 | -15.9383 |  | 2.1 | 1.79 | 20 | 38159530
21 | 0.9 | 1.11 | 0.21 | 23.3333 | 20.8955 | 0 | 1.435 | -9.0592 | 16 | 1.005 | 20.8955 |  | 0.9 | 1.11 | 21 | 32023967
23 | 0.5 | 0.72 | 0.22 | 44 | 36.0656 | 0 | 1.485 | -15.4882 | 17 | 0.61 | 36.0656 |  | 0.5 | 0.72 | 23 | 38160005
26 | 2.8 | 1.87 | -0.93 | -33.2143 | -39.8287 | 0 | 1.53 | -22.2222 | 18 | 2.335 | -39.8287 |  | 2.8 | 1.87 | 26 | 37250351
28 | 2.4 | 1.53 | -0.87 | -36.25 | -44.2748 | 0 | 1.6 | -25 | 19 | 1.965 | -44.2748 |  | 2.4 | 1.53 | 28 | 37251452
29 | 1.5 | 1.37 | -0.13 | -8.6667 | -9.0592 | 0 | 1.885 | -12.2016 | 20 | 1.435 | -9.0592 |  | 1.5 | 1.37 | 29 | 37131730
32 | 2 | 1.77 | -0.23 | -11.5 | -12.2016 | 0 | 1.92 | 2.0833 | 21 | 1.885 | -12.2016 |  | 2 | 1.77 | 32 | 37252031
34 | 3.6 | 2.55 | -1.05 | -29.1667 | -34.1463 | 0 | 1.945 | -15.9383 | 22 | 3.075 | -34.1463 |  | 3.6 | 2.55 | 34 | 37221655
37 | 0.6 | 0.63 | 0.03 | 5 | 4.878 | 0 | 1.965 | -44.2748 | 23 | 0.615 | 4.878 |  | 0.6 | 0.63 | 37 | 32127984
38 | 1.7 | 1.36 | -0.34 | -20 | -22.2222 | 0 | 2.01 | -18.9055 | 24 | 1.53 | -22.2222 |  | 1.7 | 1.36 | 38 | 38160622
39 | 2.7 | 2.06 | -0.64 | -23.7037 | -26.8908 | 0 | 2.095 | -19.5704 | 25 | 2.38 | -26.8908 |  | 2.7 | 2.06 | 39 | 32127910
41 | 6.3 | 3.48 | -2.82 | -44.7619 | -57.6687 | 0 | 2.335 | -39.8287 | 26 | 4.89 | -57.6687 |  | 6.3 | 3.48 | 41 | 37251541
45 | 9.1 | 5.2 | -3.9 | -42.8571 | -54.5455 | 0 | 2.38 | -26.8908 | 27 | 7.15 | -54.5455 |  | 9.1 | 5.2 | 45 | 37257769
46 | 5.1 | 4.27 | -0.83 | -16.2745 | -17.7161 | 0 | 2.39 | -9.205 | 28 | 4.685 | -17.7161 |  | 5.1 | 4.27 | 46 | 32143525
50 | 4.19 | 2.79 | -1.4 | -33.4129 | -40.1146 | 0 | 2.42 | -14.876 | 29 | 3.49 | -40.1146 |  | 4.19 | 2.79 | 50 | 37235534
51 | 1.2 | 1.22 | 0.02 | 1.6667 | 1.6529 | 0 | 2.58 | -40.3101 | 30 | 1.21 | 1.6529 |  | 1.2 | 1.22 | 51 | 38163096
52 | 3.1 | 2.06 | -1.04 | -33.5484 | -40.3101 | 0 | 2.69 | -23.0483 | 31 | 2.58 | -40.3101 |  | 3.1 | 2.06 | 52 | 37253621
53 | 4.1 | 3.46 | -0.64 | -15.6098 | -16.9312 | 0 | 2.875 | -1.7391 | 32 | 3.78 | -16.9312 |  | 4.1 | 3.46 | 53 | 32127967
55 | 3.7 | 2.92 | -0.78 | -21.0811 | -23.565 | 0 | 2.905 | -20.3098 | 33 | 3.31 | -23.565 |  | 3.7 | 2.92 | 55 | 37258030
57 | 1.4 | 1.44 | 0.04 | 2.8571 | 2.8169 | 0 | 3.06 | -22.2222 | 34 | 1.42 | 2.8169 |  | 1.4 | 1.44 | 57 | 37236755
58 | 1.2 | 1.07 | -0.13 | -10.8333 | -11.4537 | 0 | 3.075 | -34.1463 | 35 | 1.135 | -11.4537 |  | 1.2 | 1.07 | 58 | 32127959
59 | 2.6 | 2.24 | -0.36 | -13.8462 | -14.876 | 0 | 3.19 | -50.7837 | 36 | 2.42 | -14.876 |  | 2.6 | 2.24 | 59 | 32023879
61 | 0.7 | 0.95 | 0.25 | 35.7143 | 30.303 | 0 | 3.2 | -18.75 | 37 | 0.825 | 30.303 |  | 0.7 | 0.95 | 61 | 37085791
64 | 3.4 | 2.72 | -0.68 | -20 | -22.2222 | 0 | 3.24 | -3.7037 | 38 | 3.06 | -22.2222 |  | 3.4 | 2.72 | 64 | 32160308
65 | 1.3 | 1.12 | -0.18 | -13.8462 | -14.876 | 0 | 3.31 | -23.565 | 39 | 1.21 | -14.876 |  | 1.3 | 1.12 | 65 | 32023876
66 | 2.5 | 2.28 | -0.22 | -8.8 | -9.205 | 0 | 3.49 | -40.1146 | 40 | 2.39 | -9.205 |  | 2.5 | 2.28 | 66 | 37265382
68 | 4 | 2.38 | -1.62 | -40.5 | -50.7837 | 0 | 3.7 | -37.8378 | 41 | 3.19 | -50.7837 |  | 4 | 2.38 | 68 | 37258748
69 | 0.8 | 0.97 | 0.17 | 21.25 | 19.209 | 0 | 3.78 | -16.9312 | 42 | 0.885 | 19.209 |  | 0.8 | 0.97 | 69 | 38166192
72 | 0.5 | 0.9 | 0.4 | 80 | 57.1429 | 0 | 4.03 | -38.2134 | 43 | 0.7 | 57.1429 |  | 0.5 | 0.9 | 72 | 32143588
73 | 2.2 | 1.82 | -0.38 | -17.2727 | -18.9055 | 0 | 4.18 | -24.8804 | 44 | 2.01 | -18.9055 |  | 2.2 | 1.82 | 73 | 38167049
74 | 7.6 | 6.01 | -1.59 | -20.9211 | -23.3652 | 0 | 4.255 | -35.0176 | 45 | 6.805 | -23.3652 |  | 7.6 | 6.01 | 74 | 37267530
76 | 8.5 | 5.57 | -2.93 | -34.4706 | -41.6489 | 0 | 4.685 | -17.7161 | 46 | 7.035 | -41.6489 |  | 8.5 | 5.57 | 76 | 32137467
79 | 6.9 | 4.27 | -2.63 | -38.1159 | -47.0904 | 0 | 4.825 | -40.4145 | 47 | 5.585 | -47.0904 |  | 6.9 | 4.27 | 79 | 37272594
82 | 5.8 | 3.85 | -1.95 | -33.6207 | -40.4145 | 0 | 4.89 | -57.6687 | 48 | 4.825 | -40.4145 |  | 5.8 | 3.85 | 82 | 38169785
84 | 3.3 | 3.18 | -0.12 | -3.6364 | -3.7037 | 0 | 5.585 | -47.0904 | 49 | 3.24 | -3.7037 |  | 3.3 | 3.18 | 84 | 32147424
85 | 1.4 | 1.46 | 0.06 | 4.2857 | 4.1958 | 0 | 6.805 | -23.3652 | 50 | 1.43 | 4.1958 |  | 1.4 | 1.46 | 85 | 37274821
87 | 2.3 | 1.89 | -0.41 | -17.8261 | -19.5704 | 0 | 7.035 | -41.6489 | 51 | 2.095 | -19.5704 |  | 2.3 | 1.89 | 87 | 37273636
91 | 3.5 | 2.9 | -0.6 | -17.1429 | -18.75 | 0 | 7.15 | -54.5455 | 52 | 3.2 | -18.75 |  | 3.5 | 2.9 | 91 | 38176192
Note on A114: ###res
Data Assessment
P/B Regression
Y = 0.613 * X + 0.479
md(95) = 0.655
N = 52, r = 0.966
Cusum test for linearity shows no significant deviation from linearity.
* indicates significant difference (rejection of null-hypothesis,
gamma = 0.05 for slope or intercept from P/B and for slope from SPC).

## Werte bis 100
Method X: CEA ng/ml |  |  | Instrument: | E 170
Method Y: CEA ng/ml |  |  | Instrument: | Centaur
78
Descriptive Statistics
| X | Y | Y-X | (Y - X)% of X
Median | 4.05 | 2.96 | -0.985 | -21.6044
Mean | 21.5204 | 13.935 | -7.5854 | -20.1564
Minimum | 0.5 | 0.63 | -58.37 | -84.1066
Maximum | 98 | 80.75 | 6.52 | 80
68% Median Distance | 8.5 | 3.565 | 3.605 | 21.4755
Standard Deviation | 30.9973 | 21.4245 | 12.9832 | 26.5038
Differences
Medians |  |  |  | -26.9136
Means |  |  |  | -35.2474
Regression and Correlation Analysis
Coefficients of Correlation: | 0.94208 | 0.9021 |  |  |  |  | 4.05 | 2.96 | -0.985 | -21.6044 |  | 3.1
|  | slope b | intercept a | lower limit | upper limit |  | 21.5204 | 13.935 | -7.5854 | -20.1564 |  | 2012-04-25 12:23
Structural Relationship Model: | [6 empty] | 0.5 | 0.63 | -58.37 | -84.1066 |  | CEA
Passing/Bablok (P/B) |  | 0.63219 | 0.44867 |  |  |  | 98 | 80.75 | 6.52 | 80
95% Confidence Region for b (P/B) |  |  |  | 0.58101 | 0.683 |  | 8.5 | 3.565 | 3.605 | 21.4755 |  | ng/ml
95% Confidence Region for a (P/B) |  |  |  | 0.32504 | 0.55944 |  | 30.9973 | 21.4245 | 12.9832 | 26.5038 |  | E 170
Std. Principal Component (SPC) |  | 0.69117 | -0.93929 |  |  |  | -26.9136 | -35.2474 | 0 | 0 |  | CEA
Linear Model: | [6 empty] | 0.94208 | 0.9021 | 0.98646 | 0
Least Squares Regression |  | 0.65114 | -0.077822 |  |  |  | 0.63219 | 1 | 0.58101 | 0.683 |  | ng/ml
Theil Regression |  | 0.61698 | 0.47368 |  |  |  | 0.44867 | 1 | 0.32504 | 0.55944 |  | Centaur
Dispersion of Residuals: | [6 empty] | 0.69117 | 1 | -0.93929 | 0 |  | 78
Passing/Bablok Regression | md(68) = | 0.766 | md(95) = | 15.431 |  |  | 0.65114 | 0 | -0.077822 | 0
Std. Principal Component |  |  | SE = | 6.4542 |  |  | 1.363 | 0 | 2.5267 | 0
[7 empty] | 0.61698 | 0 | 0.47368 | 0 |  | 0.63219
[7 empty] | 0.766 | 15.431 | 6.4542 | 0 |  | Yes
[7 empty] | 0 | 0 | 1 | 78 |  | 0.44867
Yes
0.766
15.431
0.94208
0.9021
1
2
1
2
TRUE
TRUE
0
100
0
100
1
1
1
1
1
1
FALSE
FALSE
0
270
54
TRUE
Statistics | Method X: CEA ng/ml | Method Y: CEA ng/ml
N | 78 | 78
Mean | 21.5204 | 13.935
Median | 4.05 | 2.96
Minimum | 0.5 | 0.63
Maximum | 98 | 80.75
Range | 97.5 | 80.12
Difference Plot
(Normalized Differences)
| Sample Values | [18 empty] | 5 | 10 | [5 empty] | 5 | 20
Serial Number | X | Y | Y - X | (Y - X)% of X | Normalized Difference (%) | [14 empty] | 10 | 15 | [5 empty] | 10 | 25
[6 empty] | Diff Data Preset |  | Diff Ref Preset
scale_min | 0 | 0 | 0 |  |  | 0.4 | 0.05 | 0 | -0.05 |  |  |  | SelectData | [6 empty] | SelectCalibs | [6 empty] | SelectQCs | [6 empty] | F1_select |  | F2_select |  | grp_select
scale_max | 100 |  | 100 |  |  | 0.6 | -0.05 | 57.1429 | -0.05 |  |  |  | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? |  | #NAME?
| 0 |  |  | 0.44867 |  |  |  | -145.1448 |  |  |  | data_count | 78 | [5 empty] | calib_count | 0 | [5 empty] | QC_count | 0 | [5 empty] | F1_count |  | F2_count |  | selFilt_count |  |  |  |  | selCrit |  |  | H | I | J | K | L | M
| 100 |  |  | 63.6675 | [8 empty] | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | filter1 |  | filter2 |  | Selected Filters | F1 | Grp Start | Legend |  | X_Vektor | Y_Vektor | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2
1 | 1.3 | 1.56 | 0.26 | 20 | 18.1818 | 0 | 0.61 | 36.0656 | 1 | 1.43 | 18.1818 |  | 1.3 | 1.56 | 1 | 37247241
2 | 1 | 0.93 | -0.07 | -7 | -7.2539 | 0 | 0.615 | 4.878 | 2 | 0.965 | -7.2539 |  | 1 | 0.93 | 2 | 37248667
3 | 93.6 | 65.02 | -28.58 | -30.5342 | -36.0358 | 0 | 0.7 | 57.1429 | 3 | 79.31 | -36.0358 |  | 93.6 | 65.02 | 3 | 37226100
4 | 4.7 | 3.66 | -1.04 | -22.1277 | -24.8804 | 0 | 0.825 | 30.303 | 4 | 4.18 | -24.8804 |  | 4.7 | 3.66 | 4 | 37248732
5 | 1.6 | 1.37 | -0.23 | -14.375 | -15.4882 | 0 | 0.885 | 19.209 | 5 | 1.485 | -15.4882 |  | 1.6 | 1.37 | 5 | 37247206
6 | 1.8 | 1.4 | -0.4 | -22.2222 | -25 | 0 | 0.915 | 25.1366 | 6 | 1.6 | -25 |  | 1.8 | 1.4 | 6 | 32127950
7 | 3 | 2.38 | -0.62 | -20.6667 | -23.0483 | 0 | 0.965 | -7.2539 | 7 | 2.69 | -23.0483 |  | 3 | 2.38 | 7 | 32127940
8 | 0.9 | 1.04 | 0.14 | 15.5556 | 14.433 | 0 | 0.97 | 14.433 | 8 | 0.97 | 14.433 |  | 0.9 | 1.04 | 8 | 32127943
9 | 3.2 | 2.61 | -0.59 | -18.4375 | -20.3098 | 0 | 1.005 | 20.8955 | 9 | 2.905 | -20.3098 |  | 3.2 | 2.61 | 9 | 37241267
10 | 1.9 | 1.94 | 0.04 | 2.1053 | 2.0833 | 0 | 1.135 | -11.4537 | 10 | 1.92 | 2.0833 |  | 1.9 | 1.94 | 10 | 37249618
11 | 5 | 3.51 | -1.49 | -29.8 | -35.0176 | 0 | 1.21 | 1.6529 | 11 | 4.255 | -35.0176 |  | 5 | 3.51 | 11 | 37249443
14 | 4.8 | 3.26 | -1.54 | -32.0833 | -38.2134 | 0 | 1.21 | -14.876 | 12 | 4.03 | -38.2134 |  | 4.8 | 3.26 | 14 | 37250325
15 | 4.4 | 3 | -1.4 | -31.8182 | -37.8378 | 0 | 1.42 | 2.8169 | 13 | 3.7 | -37.8378 |  | 4.4 | 3 | 15 | 37250153
16 | 13.7 | 9.77 | -3.93 | -28.6861 | -33.4896 | 0 | 1.43 | 18.1818 | 14 | 11.735 | -33.4896 |  | 13.7 | 9.77 | 16 | 37242476
18 | 2.9 | 2.85 | -0.05 | -1.7241 | -1.7391 | 0 | 1.43 | 4.1958 | 15 | 2.875 | -1.7391 |  | 2.9 | 2.85 | 18 | 32161366
19 | 0.8 | 1.03 | 0.23 | 28.75 | 25.1366 | 0 | 1.435 | -9.0592 | 16 | 0.915 | 25.1366 |  | 0.8 | 1.03 | 19 | 37250080
20 | 2.1 | 1.79 | -0.31 | -14.7619 | -15.9383 | 0 | 1.485 | -15.4882 | 17 | 1.945 | -15.9383 |  | 2.1 | 1.79 | 20 | 38159530
21 | 0.9 | 1.11 | 0.21 | 23.3333 | 20.8955 | 0 | 1.53 | -22.2222 | 18 | 1.005 | 20.8955 |  | 0.9 | 1.11 | 21 | 32023967
23 | 0.5 | 0.72 | 0.22 | 44 | 36.0656 | 0 | 1.6 | -25 | 19 | 0.61 | 36.0656 |  | 0.5 | 0.72 | 23 | 38160005
24 | 11.4 | 6.15 | -5.25 | -46.0526 | -59.8291 | 0 | 1.885 | -12.2016 | 20 | 8.775 | -59.8291 |  | 11.4 | 6.15 | 24 | 37218607
26 | 2.8 | 1.87 | -0.93 | -33.2143 | -39.8287 | 0 | 1.92 | 2.0833 | 21 | 2.335 | -39.8287 |  | 2.8 | 1.87 | 26 | 37250351
27 | 25.4 | 13.18 | -12.22 | -48.1102 | -63.3489 | 0 | 1.945 | -15.9383 | 22 | 19.29 | -63.3489 |  | 25.4 | 13.18 | 27 | 37249620
28 | 2.4 | 1.53 | -0.87 | -36.25 | -44.2748 | 0 | 1.965 | -44.2748 | 23 | 1.965 | -44.2748 |  | 2.4 | 1.53 | 28 | 37251452
29 | 1.5 | 1.37 | -0.13 | -8.6667 | -9.0592 | 0 | 2.01 | -18.9055 | 24 | 1.435 | -9.0592 |  | 1.5 | 1.37 | 29 | 37131730
31 | 85.5 | 60.75 | -24.75 | -28.9474 | -33.8462 | 0 | 2.095 | -19.5704 | 25 | 73.125 | -33.8462 |  | 85.5 | 60.75 | 31 | 37249589
32 | 2 | 1.77 | -0.23 | -11.5 | -12.2016 | 0 | 2.335 | -39.8287 | 26 | 1.885 | -12.2016 |  | 2 | 1.77 | 32 | 37252031
33 | 87 | 75.26 | -11.74 | -13.4943 | -14.4706 | 0 | 2.38 | -26.8908 | 27 | 81.13 | -14.4706 |  | 87 | 75.26 | 33 | 37251102
34 | 3.6 | 2.55 | -1.05 | -29.1667 | -34.1463 | 0 | 2.39 | -9.205 | 28 | 3.075 | -34.1463 |  | 3.6 | 2.55 | 34 | 37221655
35 | 52.7 | 29.74 | -22.96 | -43.5674 | -55.7011 | 0 | 2.42 | -14.876 | 29 | 41.22 | -55.7011 |  | 52.7 | 29.74 | 35 | 37229749
37 | 0.6 | 0.63 | 0.03 | 5 | 4.878 | 0 | 2.58 | -40.3101 | 30 | 0.615 | 4.878 |  | 0.6 | 0.63 | 37 | 32127984
38 | 1.7 | 1.36 | -0.34 | -20 | -22.2222 | 0 | 2.69 | -23.0483 | 31 | 1.53 | -22.2222 |  | 1.7 | 1.36 | 38 | 38160622
39 | 2.7 | 2.06 | -0.64 | -23.7037 | -26.8908 | 0 | 2.875 | -1.7391 | 32 | 2.38 | -26.8908 |  | 2.7 | 2.06 | 39 | 32127910
40 | 23 | 15.85 | -7.15 | -31.087 | -36.8082 | 0 | 2.905 | -20.3098 | 33 | 19.425 | -36.8082 |  | 23 | 15.85 | 40 | 37252285
41 | 6.3 | 3.48 | -2.82 | -44.7619 | -57.6687 | 0 | 3.06 | -22.2222 | 34 | 4.89 | -57.6687 |  | 6.3 | 3.48 | 41 | 37251541
43 | 69.4 | 11.03 | -58.37 | -84.1066 | -145.1448 | 0 | 3.075 | -34.1463 | 35 | 40.215 | -145.1448 |  | 69.4 | 11.03 | 43 | 37253194
45 | 9.1 | 5.2 | -3.9 | -42.8571 | -54.5455 | 0 | 3.19 | -50.7837 | 36 | 7.15 | -54.5455 |  | 9.1 | 5.2 | 45 | 37257769
46 | 5.1 | 4.27 | -0.83 | -16.2745 | -17.7161 | 0 | 3.2 | -18.75 | 37 | 4.685 | -17.7161 |  | 5.1 | 4.27 | 46 | 32143525
48 | 28.9 | 35.42 | 6.52 | 22.5606 | 20.2736 | 0 | 3.24 | -3.7037 | 38 | 32.16 | 20.2736 |  | 28.9 | 35.42 | 48 | 37257518
50 | 4.19 | 2.79 | -1.4 | -33.4129 | -40.1146 | 0 | 3.31 | -23.565 | 39 | 3.49 | -40.1146 |  | 4.19 | 2.79 | 50 | 37235534
51 | 1.2 | 1.22 | 0.02 | 1.6667 | 1.6529 | 0 | 3.49 | -40.1146 | 40 | 1.21 | 1.6529 |  | 1.2 | 1.22 | 51 | 38163096
52 | 3.1 | 2.06 | -1.04 | -33.5484 | -40.3101 | 0 | 3.7 | -37.8378 | 41 | 2.58 | -40.3101 |  | 3.1 | 2.06 | 52 | 37253621
53 | 4.1 | 3.46 | -0.64 | -15.6098 | -16.9312 | 0 | 3.78 | -16.9312 | 42 | 3.78 | -16.9312 |  | 4.1 | 3.46 | 53 | 32127967
54 | 34.8 | 7.95 | -26.85 | -77.1552 | -125.614 | 0 | 4.03 | -38.2134 | 43 | 21.375 | -125.614 |  | 34.8 | 7.95 | 54 | 37232405
55 | 3.7 | 2.92 | -0.78 | -21.0811 | -23.565 | 0 | 4.18 | -24.8804 | 44 | 3.31 | -23.565 |  | 3.7 | 2.92 | 55 | 37258030
56 | 40.4 | 31.25 | -9.15 | -22.6485 | -25.5408 | 0 | 4.255 | -35.0176 | 45 | 35.825 | -25.5408 |  | 40.4 | 31.25 | 56 | 37251561
57 | 1.4 | 1.44 | 0.04 | 2.8571 | 2.8169 | 0 | 4.685 | -17.7161 | 46 | 1.42 | 2.8169 |  | 1.4 | 1.44 | 57 | 37236755
58 | 1.2 | 1.07 | -0.13 | -10.8333 | -11.4537 | 0 | 4.825 | -40.4145 | 47 | 1.135 | -11.4537 |  | 1.2 | 1.07 | 58 | 32127959
59 | 2.6 | 2.24 | -0.36 | -13.8462 | -14.876 | 0 | 4.89 | -57.6687 | 48 | 2.42 | -14.876 |  | 2.6 | 2.24 | 59 | 32023879
61 | 0.7 | 0.95 | 0.25 | 35.7143 | 30.303 | 0 | 5.585 | -47.0904 | 49 | 0.825 | 30.303 |  | 0.7 | 0.95 | 61 | 37085791
64 | 3.4 | 2.72 | -0.68 | -20 | -22.2222 | 0 | 6.805 | -23.3652 | 50 | 3.06 | -22.2222 |  | 3.4 | 2.72 | 64 | 32160308
65 | 1.3 | 1.12 | -0.18 | -13.8462 | -14.876 | 0 | 7.035 | -41.6489 | 51 | 1.21 | -14.876 |  | 1.3 | 1.12 | 65 | 32023876
66 | 2.5 | 2.28 | -0.22 | -8.8 | -9.205 | 0 | 7.15 | -54.5455 | 52 | 2.39 | -9.205 |  | 2.5 | 2.28 | 66 | 37265382
67 | 96.1 | 59.61 | -36.49 | -37.9709 | -46.8692 | 0 | 8.775 | -59.8291 | 53 | 77.855 | -46.8692 |  | 96.1 | 59.61 | 67 | 37242762
68 | 4 | 2.38 | -1.62 | -40.5 | -50.7837 | 0 | 11.35 | -78.4141 | 54 | 3.19 | -50.7837 |  | 4 | 2.38 | 68 | 37258748
69 | 0.8 | 0.97 | 0.17 | 21.25 | 19.209 | 0 | 11.735 | -33.4896 | 55 | 0.885 | 19.209 |  | 0.8 | 0.97 | 69 | 38166192
70 | 82.5 | 59.92 | -22.58 | -27.3697 | -31.709 | 0 | 19.29 | -63.3489 | 56 | 71.21 | -31.709 |  | 82.5 | 59.92 | 70 | 37259938
72 | 0.5 | 0.9 | 0.4 | 80 | 57.1429 | 0 | 19.425 | -36.8082 | 57 | 0.7 | 57.1429 |  | 0.5 | 0.9 | 72 | 32143588
73 | 2.2 | 1.82 | -0.38 | -17.2727 | -18.9055 | 0 | 21.375 | -125.614 | 58 | 2.01 | -18.9055 |  | 2.2 | 1.82 | 73 | 38167049
74 | 7.6 | 6.01 | -1.59 | -20.9211 | -23.3652 | 0 | 32.16 | 20.2736 | 59 | 6.805 | -23.3652 |  | 7.6 | 6.01 | 74 | 37267530
75 | 92.3 | 80.75 | -11.55 | -12.5135 | -13.3487 | 0 | 33.925 | -55.269 | 60 | 86.525 | -13.3487 |  | 92.3 | 80.75 | 75 | 37265007
76 | 8.5 | 5.57 | -2.93 | -34.4706 | -41.6489 | 0 | 35.825 | -25.5408 | 61 | 7.035 | -41.6489 |  | 8.5 | 5.57 | 76 | 32137467
78 | 96.4 | 76.43 | -19.97 | -20.7158 | -23.1094 | 0 | 38.345 | -36.2759 | 62 | 86.415 | -23.1094 |  | 96.4 | 76.43 | 78 | 37268394
79 | 6.9 | 4.27 | -2.63 | -38.1159 | -47.0904 | 0 | 40.215 | -145.1448 | 63 | 5.585 | -47.0904 |  | 6.9 | 4.27 | 79 | 37272594
82 | 5.8 | 3.85 | -1.95 | -33.6207 | -40.4145 | 0 | 40.77 | -130.1447 | 64 | 4.825 | -40.4145 |  | 5.8 | 3.85 | 82 | 38169785
84 | 3.3 | 3.18 | -0.12 | -3.6364 | -3.7037 | 0 | 41.22 | -55.7011 | 65 | 3.24 | -3.7037 |  | 3.3 | 3.18 | 84 | 32147424
85 | 1.4 | 1.46 | 0.06 | 4.2857 | 4.1958 | 0 | 43.37 | -70.694 | 66 | 1.43 | 4.1958 |  | 1.4 | 1.46 | 85 | 37274821
86 | 77.3 | 58.34 | -18.96 | -24.5278 | -27.9564 | 0 | 48.005 | -42.4747 | 67 | 67.82 | -27.9564 |  | 77.3 | 58.34 | 86 | 37272645
87 | 2.3 | 1.89 | -0.41 | -17.8261 | -19.5704 | 0 | 48.455 | -51.3672 | 68 | 2.095 | -19.5704 |  | 2.3 | 1.89 | 87 | 37273636
88 | 15.8 | 6.9 | -8.9 | -56.3291 | -78.4141 | 0 | 50.75 | -59.7044 | 69 | 11.35 | -78.4141 |  | 15.8 | 6.9 | 88 | 38170230
90 | 67.3 | 14.24 | -53.06 | -78.841 | -130.1447 | 0 | 67.82 | -27.9564 | 70 | 40.77 | -130.1447 |  | 67.3 | 14.24 | 90 | 37286656
91 | 3.5 | 2.9 | -0.6 | -17.1429 | -18.75 | 0 | 71.21 | -31.709 | 71 | 3.2 | -18.75 |  | 3.5 | 2.9 | 91 | 38176192
93 | 58.7 | 28.04 | -30.66 | -52.2317 | -70.694 | 0 | 73.125 | -33.8462 | 72 | 43.37 | -70.694 |  | 58.7 | 28.04 | 93 | 37285502
94 | 58.2 | 37.81 | -20.39 | -35.0344 | -42.4747 | 0 | 77.125 | -54.1329 | 73 | 48.005 | -42.4747 |  | 58.2 | 37.81 | 94
95 | 43.3 | 24.55 | -18.75 | -43.3025 | -55.269 | 0 | 77.855 | -46.8692 | 74 | 33.925 | -55.269 |  | 43.3 | 24.55 | 95
96 | 65.9 | 35.6 | -30.3 | -45.9788 | -59.7044 | 0 | 79.31 | -36.0358 | 75 | 50.75 | -59.7044 |  | 65.9 | 35.6 | 96
97 | 60.9 | 36.01 | -24.89 | -40.8703 | -51.3672 | 0 | 81.13 | -14.4706 | 76 | 48.455 | -51.3672 |  | 60.9 | 36.01 | 97
98 | 98 | 56.25 | -41.75 | -42.602 | -54.1329 | 0 | 86.415 | -23.1094 | 77 | 77.125 | -54.1329 |  | 98 | 56.25 | 98
100 | 45.3 | 31.39 | -13.91 | -30.7064 | -36.2759 | 0 | 86.525 | -13.3487 | 78 | 38.345 | -36.2759 |  | 45.3 | 31.39 | 100
Note on A114: ###res
Data Assessment
P/B Regression
Y = 0.632 * X + 0.449
md(95) = 15.431
N = 78, r = 0.942
Cusum test for linearity shows no significant deviation from linearity.
* indicates significant difference (rejection of null-hypothesis,
gamma = 0.05 for slope or intercept from P/B and for slope from SPC).

## alle Werte
Method X: CEA ng/ml |  |  | Instrument: | E 170
Method Y: CEA ng/ml |  |  | Instrument: | Centaur
90
Descriptive Statistics
| X | Y | Y-X | (Y - X)% of X
Median | 5.05 | 3.585 | -1.445 | -22.4354
Mean | 41.741 | 25.6661 | -16.0749 | -23.0063
Minimum | 0.5 | 0.63 | -165.29 | -84.1066
Maximum | 256.9 | 269.7 | 28.4 | 80
68% Median Distance | 39.25 | 21.525 | 10.535 | 22.935
Standard Deviation | 62.908 | 44.6971 | 33.6 | 27.7513
Differences
Medians |  |  |  | -29.0099
Means |  |  |  | -38.511
Regression and Correlation Analysis
Coefficients of Correlation: | 0.85822 | 0.89363 |  |  |  |  | 5.05 | 3.585 | -1.445 | -22.4354 |  | 3.1
|  | slope b | intercept a | lower limit | upper limit |  | 41.741 | 25.6661 | -16.0749 | -23.0063 |  | 2012-04-25 12:22
Structural Relationship Model: | [6 empty] | 0.5 | 0.63 | -165.29 | -84.1066 |  | CEA
Passing/Bablok (P/B) |  | 0.61564 | 0.4698 |  |  |  | 256.9 | 269.7 | 28.4 | 80
95% Confidence Region for b (P/B) |  |  |  | 0.56954 | 0.66972 |  | 39.25 | 21.525 | 10.535 | 22.935 |  | ng/ml
95% Confidence Region for a (P/B) |  |  |  | 0.34812 | 0.57722 |  | 62.908 | 44.6971 | 33.6 | 27.7513 |  | E 170
Std. Principal Component (SPC) |  | 0.71052 | -3.9915 |  |  |  | -29.0099 | -38.511 | 0 | 0 |  | CEA
Linear Model: | [6 empty] | 0.85822 | 0.89363 | 0.98231 | 0
Least Squares Regression |  | 0.60978 | 0.21338 |  |  |  | 0.61564 | 1 | 0.56954 | 0.66972 |  | ng/ml
Theil Regression |  | 0.60244 | 0.49744 |  |  |  | 0.4698 | 1 | 0.34812 | 0.57722 |  | Centaur
Dispersion of Residuals: | [6 empty] | 0.71052 | 1 | -3.9915 | 0 |  | 90
Passing/Bablok Regression | md(68) = | 2.6423 | md(95) = | 49.5513 |  |  | 0.60978 | 0 | 0.21338 | 0
Std. Principal Component |  |  | SE = | 20.6632 |  |  | 1.2079 | 0 | 10.7394 | 0
[7 empty] | 0.60244 | 0 | 0.49744 | 0 |  | 0.61564
[7 empty] | 2.6423 | 49.5513 | 20.6632 | 0 |  | Yes
[7 empty] | 0 | 0 | 1 | 90 |  | 0.4698
Yes
2.6423
49.5513
0.85822
0.89363
1
2
1
2
TRUE
TRUE
1
1
1
1
1
1
FALSE
FALSE
0
10
2
TRUE
Statistics | Method X: CEA ng/ml | Method Y: CEA ng/ml
N | 90 | 90
Mean | 41.741 | 25.6661
Median | 5.05 | 3.585
Minimum | 0.5 | 0.63
Maximum | 256.9 | 269.7
Range | 256.4 | 269.07
Difference Plot
(Normalized Differences)
| Sample Values | [18 empty] | 5 | 10 | [5 empty] | 5 | 20
Serial Number | X | Y | Y - X | (Y - X)% of X | Normalized Difference (%) | [14 empty] | 10 | 15 | [5 empty] | 10 | 25
[6 empty] | Diff Data Preset |  | Diff Ref Preset
scale_min | 0 | 0 | 0 |  |  | 0.4 | 0.05 | 0 | -0.05 |  |  |  | SelectData | [6 empty] | SelectCalibs | [6 empty] | SelectQCs | [6 empty] | F1_select |  | F2_select |  | grp_select
scale_max | 270 |  | 270 |  |  | 0.6 | -0.05 | 57.1429 | -0.05 |  |  |  | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? |  | #NAME?
| 0 |  |  | 0.4698 |  |  |  | -145.1448 |  |  |  | data_count | 90 | [5 empty] | calib_count | 0 | [5 empty] | QC_count | 0 | [5 empty] | F1_count |  | F2_count |  | selFilt_count |  |  |  |  | selCrit |  |  | H | I | J | K | L | M
| 270 |  |  | 166.6939 | [8 empty] | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | filter1 |  | filter2 |  | Selected Filters | F1 | Grp Start | Legend |  | X_Vektor | Y_Vektor | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2
1 | 1.3 | 1.56 | 0.26 | 20 | 18.1818 | 0 | 0.61 | 36.0656 | 1 | 1.43 | 18.1818 |  | 1.3 | 1.56 | 1 | 37247241
2 | 1 | 0.93 | -0.07 | -7 | -7.2539 | 0 | 0.615 | 4.878 | 2 | 0.965 | -7.2539 |  | 1 | 0.93 | 2 | 37248667
3 | 93.6 | 65.02 | -28.58 | -30.5342 | -36.0358 | 0 | 0.7 | 57.1429 | 3 | 79.31 | -36.0358 |  | 93.6 | 65.02 | 3 | 37226100
4 | 4.7 | 3.66 | -1.04 | -22.1277 | -24.8804 | 0 | 0.825 | 30.303 | 4 | 4.18 | -24.8804 |  | 4.7 | 3.66 | 4 | 37248732
5 | 1.6 | 1.37 | -0.23 | -14.375 | -15.4882 | 0 | 0.885 | 19.209 | 5 | 1.485 | -15.4882 |  | 1.6 | 1.37 | 5 | 37247206
6 | 1.8 | 1.4 | -0.4 | -22.2222 | -25 | 0 | 0.915 | 25.1366 | 6 | 1.6 | -25 |  | 1.8 | 1.4 | 6 | 32127950
7 | 3 | 2.38 | -0.62 | -20.6667 | -23.0483 | 0 | 0.965 | -7.2539 | 7 | 2.69 | -23.0483 |  | 3 | 2.38 | 7 | 32127940
8 | 0.9 | 1.04 | 0.14 | 15.5556 | 14.433 | 0 | 0.97 | 14.433 | 8 | 0.97 | 14.433 |  | 0.9 | 1.04 | 8 | 32127943
9 | 3.2 | 2.61 | -0.59 | -18.4375 | -20.3098 | 0 | 1.005 | 20.8955 | 9 | 2.905 | -20.3098 |  | 3.2 | 2.61 | 9 | 37241267
10 | 1.9 | 1.94 | 0.04 | 2.1053 | 2.0833 | 0 | 1.135 | -11.4537 | 10 | 1.92 | 2.0833 |  | 1.9 | 1.94 | 10 | 37249618
11 | 5 | 3.51 | -1.49 | -29.8 | -35.0176 | 0 | 1.21 | 1.6529 | 11 | 4.255 | -35.0176 |  | 5 | 3.51 | 11 | 37249443
12 | 241.3 | 269.7 | 28.4 | 11.7696 | 11.1155 | 0 | 1.21 | -14.876 | 12 | 255.5 | 11.1155 |  | 241.3 | 269.7 | 12 | 37248600
14 | 4.8 | 3.26 | -1.54 | -32.0833 | -38.2134 | 0 | 1.42 | 2.8169 | 13 | 4.03 | -38.2134 |  | 4.8 | 3.26 | 14 | 37250325
15 | 4.4 | 3 | -1.4 | -31.8182 | -37.8378 | 0 | 1.43 | 18.1818 | 14 | 3.7 | -37.8378 |  | 4.4 | 3 | 15 | 37250153
16 | 13.7 | 9.77 | -3.93 | -28.6861 | -33.4896 | 0 | 1.43 | 4.1958 | 15 | 11.735 | -33.4896 |  | 13.7 | 9.77 | 16 | 37242476
18 | 2.9 | 2.85 | -0.05 | -1.7241 | -1.7391 | 0 | 1.435 | -9.0592 | 16 | 2.875 | -1.7391 |  | 2.9 | 2.85 | 18 | 32161366
19 | 0.8 | 1.03 | 0.23 | 28.75 | 25.1366 | 0 | 1.485 | -15.4882 | 17 | 0.915 | 25.1366 |  | 0.8 | 1.03 | 19 | 37250080
20 | 2.1 | 1.79 | -0.31 | -14.7619 | -15.9383 | 0 | 1.53 | -22.2222 | 18 | 1.945 | -15.9383 |  | 2.1 | 1.79 | 20 | 38159530
21 | 0.9 | 1.11 | 0.21 | 23.3333 | 20.8955 | 0 | 1.6 | -25 | 19 | 1.005 | 20.8955 |  | 0.9 | 1.11 | 21 | 32023967
22 | 112 | 73.13 | -38.87 | -34.7054 | -41.9921 | 0 | 1.885 | -12.2016 | 20 | 92.565 | -41.9921 |  | 112 | 73.13 | 22 | 37221545
23 | 0.5 | 0.72 | 0.22 | 44 | 36.0656 | 0 | 1.92 | 2.0833 | 21 | 0.61 | 36.0656 |  | 0.5 | 0.72 | 23 | 38160005
24 | 11.4 | 6.15 | -5.25 | -46.0526 | -59.8291 | 0 | 1.945 | -15.9383 | 22 | 8.775 | -59.8291 |  | 11.4 | 6.15 | 24 | 37218607
25 | 256.9 | 245.22 | -11.68 | -4.5465 | -4.6523 | 0 | 1.965 | -44.2748 | 23 | 251.06 | -4.6523 |  | 256.9 | 245.22 | 25 | 37251080
26 | 2.8 | 1.87 | -0.93 | -33.2143 | -39.8287 | 0 | 2.01 | -18.9055 | 24 | 2.335 | -39.8287 |  | 2.8 | 1.87 | 26 | 37250351
27 | 25.4 | 13.18 | -12.22 | -48.1102 | -63.3489 | 0 | 2.095 | -19.5704 | 25 | 19.29 | -63.3489 |  | 25.4 | 13.18 | 27 | 37249620
28 | 2.4 | 1.53 | -0.87 | -36.25 | -44.2748 | 0 | 2.335 | -39.8287 | 26 | 1.965 | -44.2748 |  | 2.4 | 1.53 | 28 | 37251452
29 | 1.5 | 1.37 | -0.13 | -8.6667 | -9.0592 | 0 | 2.38 | -26.8908 | 27 | 1.435 | -9.0592 |  | 1.5 | 1.37 | 29 | 37131730
30 | 118 | 25.67 | -92.33 | -78.2458 | -128.5307 | 0 | 2.39 | -9.205 | 28 | 71.835 | -128.5307 |  | 118 | 25.67 | 30 | 37249575
31 | 85.5 | 60.75 | -24.75 | -28.9474 | -33.8462 | 0 | 2.42 | -14.876 | 29 | 73.125 | -33.8462 |  | 85.5 | 60.75 | 31 | 37249589
32 | 2 | 1.77 | -0.23 | -11.5 | -12.2016 | 0 | 2.58 | -40.3101 | 30 | 1.885 | -12.2016 |  | 2 | 1.77 | 32 | 37252031
33 | 87 | 75.26 | -11.74 | -13.4943 | -14.4706 | 0 | 2.69 | -23.0483 | 31 | 81.13 | -14.4706 |  | 87 | 75.26 | 33 | 37251102
34 | 3.6 | 2.55 | -1.05 | -29.1667 | -34.1463 | 0 | 2.875 | -1.7391 | 32 | 3.075 | -34.1463 |  | 3.6 | 2.55 | 34 | 37221655
35 | 52.7 | 29.74 | -22.96 | -43.5674 | -55.7011 | 0 | 2.905 | -20.3098 | 33 | 41.22 | -55.7011 |  | 52.7 | 29.74 | 35 | 37229749
36 | 105.5 | 87.82 | -17.68 | -16.7583 | -18.2909 | 0 | 3.06 | -22.2222 | 34 | 96.66 | -18.2909 |  | 105.5 | 87.82 | 36 | 32147385
37 | 0.6 | 0.63 | 0.03 | 5 | 4.878 | 0 | 3.075 | -34.1463 | 35 | 0.615 | 4.878 |  | 0.6 | 0.63 | 37 | 32127984
38 | 1.7 | 1.36 | -0.34 | -20 | -22.2222 | 0 | 3.19 | -50.7837 | 36 | 1.53 | -22.2222 |  | 1.7 | 1.36 | 38 | 38160622
39 | 2.7 | 2.06 | -0.64 | -23.7037 | -26.8908 | 0 | 3.2 | -18.75 | 37 | 2.38 | -26.8908 |  | 2.7 | 2.06 | 39 | 32127910
40 | 23 | 15.85 | -7.15 | -31.087 | -36.8082 | 0 | 3.24 | -3.7037 | 38 | 19.425 | -36.8082 |  | 23 | 15.85 | 40 | 37252285
41 | 6.3 | 3.48 | -2.82 | -44.7619 | -57.6687 | 0 | 3.31 | -23.565 | 39 | 4.89 | -57.6687 |  | 6.3 | 3.48 | 41 | 37251541
43 | 69.4 | 11.03 | -58.37 | -84.1066 | -145.1448 | 0 | 3.49 | -40.1146 | 40 | 40.215 | -145.1448 |  | 69.4 | 11.03 | 43 | 37253194
44 | 121.1 | 71.79 | -49.31 | -40.7184 | -51.1276 | 0 | 3.7 | -37.8378 | 41 | 96.445 | -51.1276 |  | 121.1 | 71.79 | 44 | 37251546
45 | 9.1 | 5.2 | -3.9 | -42.8571 | -54.5455 | 0 | 3.78 | -16.9312 | 42 | 7.15 | -54.5455 |  | 9.1 | 5.2 | 45 | 37257769
46 | 5.1 | 4.27 | -0.83 | -16.2745 | -17.7161 | 0 | 4.03 | -38.2134 | 43 | 4.685 | -17.7161 |  | 5.1 | 4.27 | 46 | 32143525
48 | 28.9 | 35.42 | 6.52 | 22.5606 | 20.2736 | 0 | 4.18 | -24.8804 | 44 | 32.16 | 20.2736 |  | 28.9 | 35.42 | 48 | 37257518
50 | 4.19 | 2.79 | -1.4 | -33.4129 | -40.1146 | 0 | 4.255 | -35.0176 | 45 | 3.49 | -40.1146 |  | 4.19 | 2.79 | 50 | 37235534
51 | 1.2 | 1.22 | 0.02 | 1.6667 | 1.6529 | 0 | 4.685 | -17.7161 | 46 | 1.21 | 1.6529 |  | 1.2 | 1.22 | 51 | 38163096
52 | 3.1 | 2.06 | -1.04 | -33.5484 | -40.3101 | 0 | 4.825 | -40.4145 | 47 | 2.58 | -40.3101 |  | 3.1 | 2.06 | 52 | 37253621
53 | 4.1 | 3.46 | -0.64 | -15.6098 | -16.9312 | 0 | 4.89 | -57.6687 | 48 | 3.78 | -16.9312 |  | 4.1 | 3.46 | 53 | 32127967
54 | 34.8 | 7.95 | -26.85 | -77.1552 | -125.614 | 0 | 5.585 | -47.0904 | 49 | 21.375 | -125.614 |  | 34.8 | 7.95 | 54 | 37232405
55 | 3.7 | 2.92 | -0.78 | -21.0811 | -23.565 | 0 | 6.805 | -23.3652 | 50 | 3.31 | -23.565 |  | 3.7 | 2.92 | 55 | 37258030
56 | 40.4 | 31.25 | -9.15 | -22.6485 | -25.5408 | 0 | 7.035 | -41.6489 | 51 | 35.825 | -25.5408 |  | 40.4 | 31.25 | 56 | 37251561
57 | 1.4 | 1.44 | 0.04 | 2.8571 | 2.8169 | 0 | 7.15 | -54.5455 | 52 | 1.42 | 2.8169 |  | 1.4 | 1.44 | 57 | 37236755
58 | 1.2 | 1.07 | -0.13 | -10.8333 | -11.4537 | 0 | 8.775 | -59.8291 | 53 | 1.135 | -11.4537 |  | 1.2 | 1.07 | 58 | 32127959
59 | 2.6 | 2.24 | -0.36 | -13.8462 | -14.876 | 0 | 11.35 | -78.4141 | 54 | 2.42 | -14.876 |  | 2.6 | 2.24 | 59 | 32023879
61 | 0.7 | 0.95 | 0.25 | 35.7143 | 30.303 | 0 | 11.735 | -33.4896 | 55 | 0.825 | 30.303 |  | 0.7 | 0.95 | 61 | 37085791
64 | 3.4 | 2.72 | -0.68 | -20 | -22.2222 | 0 | 19.29 | -63.3489 | 56 | 3.06 | -22.2222 |  | 3.4 | 2.72 | 64 | 32160308
65 | 1.3 | 1.12 | -0.18 | -13.8462 | -14.876 | 0 | 19.425 | -36.8082 | 57 | 1.21 | -14.876 |  | 1.3 | 1.12 | 65 | 32023876
66 | 2.5 | 2.28 | -0.22 | -8.8 | -9.205 | 0 | 21.375 | -125.614 | 58 | 2.39 | -9.205 |  | 2.5 | 2.28 | 66 | 37265382
67 | 96.1 | 59.61 | -36.49 | -37.9709 | -46.8692 | 0 | 32.16 | 20.2736 | 59 | 77.855 | -46.8692 |  | 96.1 | 59.61 | 67 | 37242762
68 | 4 | 2.38 | -1.62 | -40.5 | -50.7837 | 0 | 33.925 | -55.269 | 60 | 3.19 | -50.7837 |  | 4 | 2.38 | 68 | 37258748
69 | 0.8 | 0.97 | 0.17 | 21.25 | 19.209 | 0 | 35.825 | -25.5408 | 61 | 0.885 | 19.209 |  | 0.8 | 0.97 | 69 | 38166192
70 | 82.5 | 59.92 | -22.58 | -27.3697 | -31.709 | 0 | 38.345 | -36.2759 | 62 | 71.21 | -31.709 |  | 82.5 | 59.92 | 70 | 37259938
71 | 219.6 | 64.51 | -155.09 | -70.6239 | -109.176 | 0 | 40.215 | -145.1448 | 63 | 142.055 | -109.176 |  | 219.6 | 64.51 | 71 | 37259923
72 | 0.5 | 0.9 | 0.4 | 80 | 57.1429 | 0 | 40.77 | -130.1447 | 64 | 0.7 | 57.1429 |  | 0.5 | 0.9 | 72 | 32143588
73 | 2.2 | 1.82 | -0.38 | -17.2727 | -18.9055 | 0 | 41.22 | -55.7011 | 65 | 2.01 | -18.9055 |  | 2.2 | 1.82 | 73 | 38167049
74 | 7.6 | 6.01 | -1.59 | -20.9211 | -23.3652 | 0 | 43.37 | -70.694 | 66 | 6.805 | -23.3652 |  | 7.6 | 6.01 | 74 | 37267530
75 | 92.3 | 80.75 | -11.55 | -12.5135 | -13.3487 | 0 | 48.005 | -42.4747 | 67 | 86.525 | -13.3487 |  | 92.3 | 80.75 | 75 | 37265007
76 | 8.5 | 5.57 | -2.93 | -34.4706 | -41.6489 | 0 | 48.455 | -51.3672 | 68 | 7.035 | -41.6489 |  | 8.5 | 5.57 | 76 | 32137467
77 | 133.3 | 63.82 | -69.48 | -52.123 | -70.4951 | 0 | 50.75 | -59.7044 | 69 | 98.56 | -70.4951 |  | 133.3 | 63.82 | 77 | 38167355
78 | 96.4 | 76.43 | -19.97 | -20.7158 | -23.1094 | 0 | 67.82 | -27.9564 | 70 | 86.415 | -23.1094 |  | 96.4 | 76.43 | 78 | 37268394
79 | 6.9 | 4.27 | -2.63 | -38.1159 | -47.0904 | 0 | 71.21 | -31.709 | 71 | 5.585 | -47.0904 |  | 6.9 | 4.27 | 79 | 37272594
81 | 201.7 | 90.25 | -111.45 | -55.2553 | -76.3487 | 0 | 71.835 | -128.5307 | 72 | 145.975 | -76.3487 |  | 201.7 | 90.25 | 81 | 37257202
82 | 5.8 | 3.85 | -1.95 | -33.6207 | -40.4145 | 0 | 73.125 | -33.8462 | 73 | 4.825 | -40.4145 |  | 5.8 | 3.85 | 82 | 38169785
83 | 221.3 | 67.78 | -153.52 | -69.3719 | -106.2128 | 0 | 77.125 | -54.1329 | 74 | 144.54 | -106.2128 |  | 221.3 | 67.78 | 83 | 37273311
84 | 3.3 | 3.18 | -0.12 | -3.6364 | -3.7037 | 0 | 77.855 | -46.8692 | 75 | 3.24 | -3.7037 |  | 3.3 | 3.18 | 84 | 32147424
85 | 1.4 | 1.46 | 0.06 | 4.2857 | 4.1958 | 0 | 79.31 | -36.0358 | 76 | 1.43 | 4.1958 |  | 1.4 | 1.46 | 85 | 37274821
86 | 77.3 | 58.34 | -18.96 | -24.5278 | -27.9564 | 0 | 81.13 | -14.4706 | 77 | 67.82 | -27.9564 |  | 77.3 | 58.34 | 86 | 37272645
87 | 2.3 | 1.89 | -0.41 | -17.8261 | -19.5704 | 0 | 86.415 | -23.1094 | 78 | 2.095 | -19.5704 |  | 2.3 | 1.89 | 87 | 37273636
88 | 15.8 | 6.9 | -8.9 | -56.3291 | -78.4141 | 0 | 86.525 | -13.3487 | 79 | 11.35 | -78.4141 |  | 15.8 | 6.9 | 88 | 38170230
89 | 117 | 98.22 | -18.78 | -16.0513 | -17.4519 | 0 | 92.565 | -41.9921 | 80 | 107.61 | -17.4519 |  | 117 | 98.22 | 89 | 37277819
90 | 67.3 | 14.24 | -53.06 | -78.841 | -130.1447 | 0 | 96.445 | -51.1276 | 81 | 40.77 | -130.1447 |  | 67.3 | 14.24 | 90 | 37286656
91 | 3.5 | 2.9 | -0.6 | -17.1429 | -18.75 | 0 | 96.66 | -18.2909 | 82 | 3.2 | -18.75 |  | 3.5 | 2.9 | 91 | 38176192
93 | 58.7 | 28.04 | -30.66 | -52.2317 | -70.694 | 0 | 98.56 | -70.4951 | 83 | 43.37 | -70.694 |  | 58.7 | 28.04 | 93 | 37285502
94 | 58.2 | 37.81 | -20.39 | -35.0344 | -42.4747 | 0 | 107.61 | -17.4519 | 84 | 48.005 | -42.4747 |  | 58.2 | 37.81 | 94
95 | 43.3 | 24.55 | -18.75 | -43.3025 | -55.269 | 0 | 142.055 | -109.176 | 85 | 33.925 | -55.269 |  | 43.3 | 24.55 | 95
96 | 65.9 | 35.6 | -30.3 | -45.9788 | -59.7044 | 0 | 144.54 | -106.2128 | 86 | 50.75 | -59.7044 |  | 65.9 | 35.6 | 96
97 | 60.9 | 36.01 | -24.89 | -40.8703 | -51.3672 | 0 | 145.975 | -76.3487 | 87 | 48.455 | -51.3672 |  | 60.9 | 36.01 | 97
98 | 98 | 56.25 | -41.75 | -42.602 | -54.1329 | 0 | 147.755 | -111.8676 | 88 | 77.125 | -54.1329 |  | 98 | 56.25 | 98
100 | 45.3 | 31.39 | -13.91 | -30.7064 | -36.2759 | 0 | 251.06 | -4.6523 | 89 | 38.345 | -36.2759 |  | 45.3 | 31.39 | 100
104 | 230.4 | 65.11 | -165.29 | -71.7405 | -111.8676 | 0 | 255.5 | 11.1155 | 90 | 147.755 | -111.8676 |  | 230.4 | 65.11 | 104
Note on A114: ###res
Data Assessment
P/B Regression
Y = 0.616 * X + 0.470
md(95) = 49.551
N = 90, r = 0.858
Cusum test for linearity shows no significant deviation from linearity.
* indicates significant difference (rejection of null-hypothesis,
gamma = 0.05 for slope or intercept from P/B and for slope from SPC).

## Result (hidden)
Method X: |  |  | Instrument:
Method Y: |  |  | Instrument:
0
Descriptive Statistics
| X | Y | Y-X | (Y - X)% of X
Median | 0 | 0 | 0 | 0
Mean | 0 | 0 | 0 | NV
Minimum | 0 | 0 | 0 | 0
Maximum | 0 | 0 | 0 | 0
68% Median Distance | 0 | 0 | 0 | 0
Standard Deviation | 0 | 0 | 0 | NV
Differences
Medians |  |  |  | 0
Means |  |  |  | 0
Regression and Correlation Analysis
Coefficients of Correlation: | 0 | 0
|  | slope b | intercept a | lower limit | upper limit
Structural Relationship Model:
Passing/Bablok (P/B) |  | 0 | 0
95% Confidence Region for b (P/B) |  |  |  | 0 | 0
95% Confidence Region for a (P/B) |  |  |  | 0 | 0
Std. Principal Component (SPC) |  | 0 | 0
Linear Model:
Least Squares Regression |  | 0 | 0
Theil Regression |  | 0 | 0
Dispersion of Residuals:
Passing/Bablok Regression | md(68) = | 0 | md(95) = | 0
Std. Principal Component |  |  | SE = | 0
Statistics | Method X: | Method Y:
N | 0 | 0
Mean | 0 | 0
Median | 0 | 0
Minimum | 0 | 0
Maximum | 0 | 0
Range | 0 | 0
Difference Plot
(Normalized Differences)
| Sample Values | [18 empty] | 5 | 10 | [5 empty] | 5 | 20
Serial Number | X | Y | Y - X | (Y - X)% of X | Normalized Difference (%) | [14 empty] | 10 | 15 | [5 empty] | 10 | 25
[6 empty] | Diff Data Preset |  | Diff Ref Preset
scale_min | 0 | 0 | 0 |  |  | 0.4 | 0.05 | 0 | -0.05 |  |  |  | SelectData | [6 empty] | SelectCalibs | [6 empty] | SelectQCs | [6 empty] | F1_select |  | F2_select |  | grp_select
scale_max | 100 |  | 80 |  |  | 0.6 | -0.05 | 1 | -0.05 |  |  |  | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? |  | #NAME?
| 0 |  |  | 5 | [7 empty] | data_count | [6 empty] | calib_count | [6 empty] | QC_count | [6 empty] | F1_count |  | F2_count |  | selFilt_count |  |  |  |  | selCrit |  |  | H | I | J | K | L | M
| 100 |  |  | 95 | [8 empty] | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | filter1 |  | filter2 |  | Selected Filters | F1 | Grp Start | Legend |  | X_Vektor | Y_Vektor | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2
1 | 24.5 | 26.2 | 1.7 | 6.9388 | 6.7061 | 0 | 0 | 0 | 1 | 25.35 | 6.7061 | [8 empty] | -5 | 10 | [5 empty] | -5 | 20
2 | [8 empty] | 2 | [10 empty] | -10 | 15 | [5 empty] | -10 | 25
Note on A114: ###res
Data Assessment