## Summary
Summary
Subject:
Laboratory:
Result | Date | Laboratory | Material |  | Method X |  |  |  | Method Y | [6 empty] | Passing / Bablok |  |  | Disp. of Residuals |  |  | Kendall's
Sheet |  |  |  | Short Name | Long Name | Unit | Instrument | Short Name | Long Name | Unit | Instrument | n | Min X | Max X | B/P b | sign. (b) | B/P a | sign. (a) | md(68) | md(95) | r | Tau | Linearity
[13 empty] | 0 | 0 | 0 |  | 0 |  | 0 | 0 | 0 | 0

## Input Sheet
|  |  |  | Worksheet for Method Comparison | [40 empty] | 20
Subject:
Lab.:
Serial No. | Identifier | F1 | F2 | Vector 1 | Vector 2 | Vector 3 | Vector 4 | Vector 5 | Vector 6 | Vector 7 | Vector 8 | Vector 9 | Vector 10 | Vector 11 | Vector 12 | Vector 13 | Vector 14 | Vector 15 | Vector 16 | Vector 17 | Vector 18 | Vector 19 | Vector 20 | Vector 21 | Vector 22 | Vector 23 | Vector 24 | Vector 25 | Vector 26 | Vector 27 | Vector 28 | Vector 29 | Vector 30 | Vector 31 | Vector 32
| Short Name |  |  | PSA | PSA
Long Name
| [Unit] |  |  | ng/ml | ng/ml
| Instrument |  |  | E 170 | Centaur
lfd_no | ident_no | filter1 | filter2 | X_Vektor | Y_Vektor | [31 empty] | sel_X_vector | sel_Y_vector | sel_seq | sel_ident | sel_f1 | sel_f2
1 | 38158289 |  |  | 18 | 15.54 | [37 empty] | SumData
2 | 38158688 |  |  | 27.4 | 25.5 | [37 empty] | 3.1 | ver_Info
3 | 38158035 |  |  | 0.166 | 0.19 | [37 empty] | 2012-04-26 10:06 | L_DateTime
4 | 38158883 |  |  | 1.08 | 1.05 | [36 empty] | Kriterienfeld | PSA | L_SName_X
5 | 37239128 |  |  | 0.63 | 0.59 | [36 empty] | #NAME? |  | L_LName_X
6 | 38158959 |  |  | 49.9 | 44.53 | [37 empty] | ng/ml | L_Unit_X
7 | 37249463 |  |  | 0.088 | 0.09 | [37 empty] | E 170 | L_Instr_X
8 | 32127996 |  |  | 0.36 | 0.34 | [31 empty] | Zielbereich | [5 empty] | PSA | L_SName_Y
9 | 38159485 |  |  | 0.308 | 0.27 | [38 empty] | L_LName_Y
10 | 32019594 |  |  | 0.091 | 0.09 | [31 empty] | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 | ng/ml | L_Unit_Y
11 | 37250625 |  |  | 1.09 | 1 | [30 empty] | sel_first_row | 0.166 | 0.19 | 3 | 38158035 |  |  | Centaur | L_Instr_Y
12 | 38159496 |  |  | 0.972 | 0.98 | [31 empty] | 1.08 | 1.05 | 4 | 38158883 |  |  |  | L_N
13 | 38158965 |  |  | 0.753 | 0.68 | [31 empty] | 0.63 | 0.59 | 5 | 37239128 |  |  |  | L_Min_X
14 | 38159754 |  |  | 0.981 | 0.81 | [31 empty] | 0.088 | 0.09 | 7 | 37249463 |  |  |  | L_Max_X
15 | 32157083 |  |  | 8.72 | 8.23 | [31 empty] | 0.36 | 0.34 | 8 | 32127996 |  |  |  | L_PBa
16 | 38158531 |  |  | 0.167 | 0.14 | [31 empty] | 0.308 | 0.27 | 9 | 38159485 |  |  |  | L_PBas
17 | 38160092 |  |  | 5 | 4.75 | [31 empty] | 0.091 | 0.09 | 10 | 32019594 |  |  |  | L_PBb
18 | 37252133 |  |  | 0.501 | 0.43 | [31 empty] | 1.09 | 1 | 11 | 37250625 |  |  |  | L_PBbs
19 | 32023925 |  |  | 0.422 | 0.33 | [31 empty] | 0.972 | 0.98 | 12 | 38159496 |  |  |  | L_md65
20 | 38160098 |  |  | 0.233 | 0.22 | [31 empty] | 0.753 | 0.68 | 13 | 38158965 |  |  |  | L_md95
21 | 32121489 |  |  | 5.22 | 4.2 | [31 empty] | 0.981 | 0.81 | 14 | 38159754 |  |  |  | L_r
22 | 37252141 |  |  | 1.58 | 1.33 | [31 empty] | 0.167 | 0.14 | 16 | 38158531 |  |  |  | L_Tau
23 | 37252658 |  |  | 0.064 | 0.08 | [31 empty] | 0.501 | 0.43 | 18 | 37252133 |  |  |  | L_linear
24 | 32160700 |  |  | 0.382 | 0.38 | [31 empty] | 0.422 | 0.33 | 19 | 32023925 |  |  | 1 | dlg_x_vector
25 | 32157144 |  |  | 50.1 | 44.52 | [31 empty] | 0.233 | 0.22 | 20 | 38160098 |  |  | 2 | dlg_y_vector
26 | 32160702 |  |  | 0.314 | 0.31 | [31 empty] | 1.58 | 1.33 | 22 | 37252141 |  |  | 1 | dlg_Task_sel
27 | 32160703 |  |  | 17.2 | 16.07 | [31 empty] | 0.064 | 0.08 | 23 | 37252658 |  |  | 2 | dlg_plot_name
28 | 38160723 |  |  | 0.633 | 0.52 | [31 empty] | 0.382 | 0.38 | 24 | 32160700 |  |  | TRUE | dlg_plot_do_unit
29 | 32160704 |  |  | 65.9 | 50.5 | [31 empty] | 0.314 | 0.31 | 26 | 32160702 |  |  | TRUE | dlg_plot_do_instr
30 | 32160724 |  |  | 41 | 32.48 | [31 empty] | 0.633 | 0.52 | 28 | 38160723 |  |  |  | Dlg_Min_ID_No
31 | 37254199 |  |  | 1.69 | 1.73 | [31 empty] | 1.69 | 1.73 | 31 | 37254199 |  |  |  | Dlg_Max_ID_No
32 | 37254857 |  |  | 0.477 | 0.46 | [31 empty] | 0.477 | 0.46 | 32 | 37254857 |  |  |  | Dlg_Min_Identification
33 | 38162438 |  |  | 0.921 | 0.51 | [31 empty] | 0.921 | 0.51 | 33 | 38162438 |  |  |  | Dlg_Max_Identification
34 | 38162540 |  |  | 0.108 | 0.11 | [31 empty] | 0.108 | 0.11 | 34 | 38162540 |  |  | 0 | Dlg_Min_X_Value
35 | 38162433 |  |  | 1.1 | 0.84 | [31 empty] | 1.1 | 0.84 | 35 | 38162433 |  |  | 2 | Dlg_Max_X_Value
36 | 37254297 |  |  | 0.218 | 0.19 | [31 empty] | 0.218 | 0.19 | 36 | 37254297 |  |  | 0 | Dlg_Min_Y_Value
37 | 38163109 |  |  | 0.003 | 0.05 | [31 empty] | 0.003 | 0.05 | 37 | 38163109 |  |  | 2 | Dlg_Max_Y_Value
38 | 38162501 |  |  | 8.38 | 8.03 | [31 empty] | 1.03 | 0.71 | 39 | 38162503 |  |  |  | Dlg_Filter_1_Selection
39 | 38162503 |  |  | 1.03 | 0.71 | [31 empty] | 0.03 | 0.06 | 42 | 32136776 |  |  |  | Dlg_Filter_2_Selection
40 | 37264171 |  |  | 4.81 | 4.69 | [31 empty] | 0.125 | 0.15 | 44 | 38165868 |  |  | 1 | Dlg_ID_No_IncExc
41 | 32160738 |  |  | 5.21 | 4 | [31 empty] | 1.72 | 1.7 | 45 | 37261409 |  |  | 1 | Dlg_Identification_IncExc
42 | 32136776 |  |  | 0.03 | 0.06 | [31 empty] | 0.343 | 0.33 | 46 | 37261865 |  |  | 1 | Dlg_X_Value_IncExc
43 | 32127959 |  |  | 8.25 | 7.67 | [31 empty] | 1.02 | 1.03 | 47 | 37261874 |  |  | 1 | Dlg_Y_Value_IncExc
44 | 38165868 |  |  | 0.125 | 0.15 | [31 empty] | 0.009 | 0.04 | 48 | 38165776 |  |  | 1 | Dlg_Filter1_IncExc
45 | 37261409 |  |  | 1.72 | 1.7 | [31 empty] | 1.65 | 1.43 | 50 | 38165773 |  |  | 1 | Dlg_Filter2_IncExc
46 | 37261865 |  |  | 0.343 | 0.33 | [31 empty] | 0.003 | 0.04 | 51 | 38165845 |  |  | FALSE | Dlg_ShowCal
47 | 37261874 |  |  | 1.02 | 1.03 | [31 empty] | 1.03 | 0.85 | 52 | 32144696 |  |  | FALSE | Dlg_ShowQC
48 | 38165776 |  |  | 0.009 | 0.04 | [31 empty] | 0.337 | 0.33 | 53 | 37264662 |  |  | 0 | dlg_scale_min
49 | 38165229 |  |  | 40.2 | 39.02 | [31 empty] | 1.66 | 1.7 | 54 | 32023880 |  |  | 1.8 | dlg_scale_max
50 | 38165773 |  |  | 1.65 | 1.43 | [31 empty] | 0.27 | 0.3 | 55 | 37253832 |  |  | 0.36 | dlg_scale_int
51 | 38165845 |  |  | 0.003 | 0.04 | [31 empty] | 0.003 | 0.02 | 56 | 32023881 |  |  | TRUE | dlg_show_grid
52 | 32144696 |  |  | 1.03 | 0.85 | [31 empty] | 1.04 | 1.04 | 57 | 38166192
53 | 37264662 |  |  | 0.337 | 0.33 | [31 empty] | 0.784 | 0.79 | 58 | 32144698
54 | 32023880 |  |  | 1.66 | 1.7 | [31 empty] | 0.634 | 0.48 | 59 | 32023882
55 | 37253832 |  |  | 0.27 | 0.3 | [31 empty] | 1.09 | 1.03 | 60 | 38166993
56 | 32023881 |  |  | 0.003 | 0.02 | [31 empty] | 0.445 | 0.46 | 61 | 32101002
57 | 38166192 |  |  | 1.04 | 1.04 | [31 empty] | 0.02 | 0.06 | 66 | 32144701
58 | 32144698 |  |  | 0.784 | 0.79 | [31 empty] | 1.4 | 1.16 | 67 | 37272594
59 | 32023882 |  |  | 0.634 | 0.48 | [31 empty] | 0.254 | 0.27 | 71 | 37272564
60 | 38166993 |  |  | 1.09 | 1.03 | [31 empty] | 1.35 | 1.38 | 78 | 32144703
61 | 32101002 |  |  | 0.445 | 0.46 | [31 empty] | 1.37 | 1.45 | 89 | 32003034
62 | 32137484 |  |  | 22.8 | 21.48
63 | 32160744 |  |  | 5.28 | 4.51
64 | 38167182 |  |  | 5.26 | 5.04
65 | 32144700 |  |  | 46.8 | 43.26
66 | 32144701 |  |  | 0.02 | 0.06
67 | 37272594 |  |  | 1.4 | 1.16
68 | 38169363 |  |  | 19.1 | 20.66
69 | 32160613 |  |  | 8.09 | 7.46
70 | 38170528 |  |  | 4.84 | 5.3
71 | 37272564 |  |  | 0.254 | 0.27
72 | 38170067 |  |  | 7.59 | 7.92
73 | 32160615 |  |  | 53.9 | 54.88
74 | 37274073 |  |  | 56.7 | 57.33
75 | 38170698 |  |  | 26.1 | 25.19
76 | 38171774 |  |  | 5.21 | 5.53
77 | 38173004 |  |  | 19.8 | 18.85
78 | 32144703 |  |  | 1.35 | 1.38
79 | 32160631 |  |  | 8.15 | 8.76
80 | 32160629 |  |  | 7.36 | 7.63
81 | 32160628 |  |  | 81.4 | 75.25
82 | 38173363 |  |  | 31.8 | 33.75
83 | 32160632 |  |  | 21.2 | 18.5
84 | 38176246 |  |  | 24.2 | 19.72
85 | 32160633 |  |  | 5.23 | 4.7
86 | 37284567 |  |  | 6.75 | 6.69
87 | 38175501 |  |  | 65.6
88 | 38175529 |  |  | 65 | 57.9
89 | 32003034 |  |  | 1.37 | 1.45
90 | 37287086 |  |  | 7.08 | 7.25
91 | 38176759 |  |  | 7.93 | 7.42
92 |  |  |  | 2.73 | 2.73
93 |  |  |  | 2.04 | 1.98
94 |  |  |  | 4.44 | 4.36
95 |  |  |  | 6.24 | 5.63
96 |  |  |  | 3.83 | 3.58
97 |  |  |  | 3.01 | 2.43
98 |  |  |  | 2.36 | 2.2
99 |  |  |  | 3.61 | 3.07
100 |  |  |  | 6.11 | 6.58
101 |  |  |  | 6.03 | 6.22
102 |  |  |  | 5.45 | 5.51
103 |  |  |  | 2.58 | 2.69
104
105
106
107
108
109
110
111
112
113
114
115
116
117
118
119
120
121
122
123
124
125
126
127
128
129
130
131
132
133
134
135
136
137
138
139
140
141
142
143
144
145
146
147
148
149
150
151
152
153
154
155
156
157
158
159
160
161
162
163
164
165
166
167
168
169
170
171
172
173
174
175
176
177
178
179
180
181
182
183
184
185
186
187
188
189
190
191
192
193
194
195
196
197
198
199
200
201
202
203
204
205
206
207
208
209
210
211
212
213
214
215
216
217
218
219
220
221
222
223
224
225
226
227
228
229
230
231
232
233
234
235
236
237
238
239
240
241
242
243
244
245
246
247
248
249
250
251
252
253
254
255
256
257
258
259
260
261
262
263
264
265
266
267
268
269
270
271
272
273
274
275
276
277
278
279
280
281
282
283
284
285
286
287
288
289
290
291
292
293
294
295
296
297
298
299
300
301
302
303
304
305
306
307
308
309
310
311
312
313
314
315
316
317
318
319
320
321
322
323
324
325
326
327
328
329
330
331
332
333
334
335
336
337
338
339
340
341
342
343
344
345
346
347
348
349
350
351
352
353
354
355
356
357
358
359
360
361
362
363
364
365
366
367
368
369
370
371
372
373
374
375
376
377
378
379
380
381
382
383
384
385
386
387
388
389
390
391
392
393
394
395
396
397
398
399
400
401
402
403
404
405
406
407
408
409
410
411
412
413
414
415
416
417
418
419
420
421
422
423
424
425
426
427
428
429
430
431
432
433
434
435
436
437
438
439
440
441
442
443
444
445
446
447
448
449
450
451
452
453
454
455
456
457
458
459
460
461
462
463
464
465
466
467
468
469
470
471
472
473
474
475
476
477
478
479
480
481
482
483
484
485
486
487
488
489
490
491
492
493
494
495
496
497
498
499
500
501
502
503
504
505
506
507
508
509
510
511
512
513
514
515
516
517
518
519
520
521
522
523
524
525
526
527
528
529
530
531
532
533
534
535
536
537
538
539
540
541
542
543
544
545
546
547
548
549
550
551
552
553
554
555
556
557
558
559
560
561
562
563
564
565
566
567
568
569
570
571
572
573
574
575
576
577
578
579
580
581
582
583
584
585
586
587
588
589
590
591
592
593
594
595
596
597
598
599
600
601

## Werte bis 2
Method X: PSA ng/ml |  |  | Instrument: | E 170
Method Y: PSA ng/ml |  |  | Instrument: | Centaur
51
Descriptive Statistics
| X | Y | Y-X | (Y - X)% of X
Median | 0.477 | 0.46 | -0.007 | -1.2739
Mean | 0.64606 | 0.59863 | -0.047431 | 73.6014
Minimum | 0.003 | 0.02 | -0.411 | -44.6254
Maximum | 1.72 | 1.73 | 0.08 | 1566.6667
68% Median Distance | 0.504 | 0.42 | 0.053 | 15.869
Standard Deviation | 0.53278 | 0.49901 | 0.1048 | 289.7534
Differences
Medians |  |  |  | -3.5639
Means |  |  |  | -7.3416
Regression and Correlation Analysis
Coefficients of Correlation: | 0.98149 | 0.92471 |  |  |  |  | 0.477 | 0.46 | -0.007 | -1.2739 |  | 3.1
|  | slope b | intercept a | lower limit | upper limit |  | 0.64606 | 0.59863 | -0.047431 | 73.6014 |  | 2012-04-26 10:06
Structural Relationship Model: | [6 empty] | 0.003 | 0.02 | -0.411 | -44.6254 |  | PSA
Passing/Bablok (P/B) |  | 0.91977 | 0.020037 |  |  |  | 1.72 | 1.73 | 0.08 | 1566.6667
95% Confidence Region for b (P/B) |  |  |  | 0.85714 | 0.97187 |  | 0.504 | 0.42 | 0.053 | 15.869 |  | ng/ml
95% Confidence Region for a (P/B) |  |  |  | 0.0024808 | 0.036 |  | 0.53278 | 0.49901 | 0.1048 | 289.7534 |  | E 170
Std. Principal Component (SPC) |  | 0.93662 | -0.0064824 |  |  |  | -3.5639 | -7.3416 | 0 | 0 |  | PSA
Linear Model: | [6 empty] | 0.98149 | 0.92471 | 0.98921 | 0
Least Squares Regression |  | 0.91928 | 0.0047192 |  |  |  | 0.91977 | 1 | 0.85714 | 0.97187 |  | ng/ml
Theil Regression |  | 0.90287 | 0.022216 |  |  |  | 0.020037 | 1 | 0.0024808 | 0.036 |  | Centaur
Dispersion of Residuals: | [6 empty] | 0.93662 | 1 | -0.0064824 | 0 |  | 51
Passing/Bablok Regression | md(68) = | 0.052844 | md(95) = | 0.13309 |  |  | 0.91928 | 0 | 0.0047192 | 0
Std. Principal Component |  |  | SE = | 0.070942 |  |  | 1.0479 | 0 | 0.018752 | 0
[7 empty] | 0.90287 | 0 | 0.022216 | 0 |  | 0.91977
[7 empty] | 0.052844 | 0.13309 | 0.070942 | 0 |  | Yes
[7 empty] | 0 | 0 | 1 | 51 |  | 0.020037
Yes
0.052844
0.13309
0.98149
0.92471
1
2
1
2
TRUE
TRUE
0
2
0
2
1
1
1
1
1
1
FALSE
FALSE
0
9
1.8
TRUE
Statistics | Method X: PSA ng/ml | Method Y: PSA ng/ml
N | 51 | 51
Mean | 0.64606 | 0.59863
Median | 0.477 | 0.46
Minimum | 0.003 | 0.02
Maximum | 1.72 | 1.73
Range | 1.717 | 1.71
Difference Plot
(Normalized Differences)
| Sample Values | [18 empty] | 5 | 10 | [5 empty] | 5 | 20
Serial Number | X | Y | Y - X | (Y - X)% of X | Normalized Difference (%) | [14 empty] | 10 | 15 | [5 empty] | 10 | 25
[6 empty] | Diff Data Preset |  | Diff Ref Preset
scale_min | 0 | 0 | 0 |  |  | 0.4 | 0.05 | 0 | -0.05 |  |  |  | SelectData | [6 empty] | SelectCalibs | [6 empty] | SelectQCs | [6 empty] | F1_select |  | F2_select |  | grp_select
scale_max | 1.8 |  | 1.8 |  |  | 0.6 | -0.05 | 177.3585 | -0.05 |  |  |  | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? |  | #NAME?
| 0 |  |  | 0.020037 |  |  |  | -57.4423 |  |  |  | data_count | 51 | [5 empty] | calib_count | 0 | [5 empty] | QC_count | 0 | [5 empty] | F1_count |  | F2_count |  | selFilt_count |  |  |  |  | selCrit |  |  | H | I | J | K | L | M
| 1.8 |  |  | 1.6756 | [8 empty] | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | filter1 |  | filter2 |  | Selected Filters | F1 | Grp Start | Legend |  | X_Vektor | Y_Vektor | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2
3 | 0.166 | 0.19 | 0.024 | 14.4578 | 13.4831 | 0 | 0.0115 | 147.8261 | 1 | 0.178 | 13.4831 |  | 0.166 | 0.19 | 3 | 38158035
4 | 1.08 | 1.05 | -0.03 | -2.7778 | -2.8169 | 0 | 0.0215 | 172.093 | 2 | 1.065 | -2.8169 |  | 1.08 | 1.05 | 4 | 38158883
5 | 0.63 | 0.59 | -0.04 | -6.3492 | -6.5574 | 0 | 0.0245 | 126.5306 | 3 | 0.61 | -6.5574 |  | 0.63 | 0.59 | 5 | 37239128
7 | 0.088 | 0.09 | 0.002 | 2.2727 | 2.2472 | 0 | 0.0265 | 177.3585 | 4 | 0.089 | 2.2472 |  | 0.088 | 0.09 | 7 | 37249463
8 | 0.36 | 0.34 | -0.02 | -5.5556 | -5.7143 | 0 | 0.04 | 100 | 5 | 0.35 | -5.7143 |  | 0.36 | 0.34 | 8 | 32127996
9 | 0.308 | 0.27 | -0.038 | -12.3377 | -13.1488 | 0 | 0.045 | 66.6667 | 6 | 0.289 | -13.1488 |  | 0.308 | 0.27 | 9 | 38159485
10 | 0.091 | 0.09 | -0.001 | -1.0989 | -1.105 | 0 | 0.072 | 22.2222 | 7 | 0.0905 | -1.105 |  | 0.091 | 0.09 | 10 | 32019594
11 | 1.09 | 1 | -0.09 | -8.2569 | -8.6124 | 0 | 0.089 | 2.2472 | 8 | 1.045 | -8.6124 |  | 1.09 | 1 | 11 | 37250625
12 | 0.972 | 0.98 | 0.008 | 0.82305 | 0.81967 | 0 | 0.0905 | -1.105 | 9 | 0.976 | 0.81967 |  | 0.972 | 0.98 | 12 | 38159496
13 | 0.753 | 0.68 | -0.073 | -9.6946 | -10.1884 | 0 | 0.109 | 1.8349 | 10 | 0.7165 | -10.1884 |  | 0.753 | 0.68 | 13 | 38158965
14 | 0.981 | 0.81 | -0.171 | -17.4312 | -19.0955 | 0 | 0.1375 | 18.1818 | 11 | 0.8955 | -19.0955 |  | 0.981 | 0.81 | 14 | 38159754
16 | 0.167 | 0.14 | -0.027 | -16.1677 | -17.5896 | 0 | 0.1535 | -17.5896 | 12 | 0.1535 | -17.5896 |  | 0.167 | 0.14 | 16 | 38158531
18 | 0.501 | 0.43 | -0.071 | -14.1717 | -15.2524 | 0 | 0.178 | 13.4831 | 13 | 0.4655 | -15.2524 |  | 0.501 | 0.43 | 18 | 37252133
19 | 0.422 | 0.33 | -0.092 | -21.8009 | -24.4681 | 0 | 0.204 | -13.7255 | 14 | 0.376 | -24.4681 |  | 0.422 | 0.33 | 19 | 32023925
20 | 0.233 | 0.22 | -0.013 | -5.5794 | -5.7395 | 0 | 0.2265 | -5.7395 | 15 | 0.2265 | -5.7395 |  | 0.233 | 0.22 | 20 | 38160098
22 | 1.58 | 1.33 | -0.25 | -15.8228 | -17.1821 | 0 | 0.262 | 6.1069 | 16 | 1.455 | -17.1821 |  | 1.58 | 1.33 | 22 | 37252141
23 | 0.064 | 0.08 | 0.016 | 25 | 22.2222 | 0 | 0.285 | 10.5263 | 17 | 0.072 | 22.2222 |  | 0.064 | 0.08 | 23 | 37252658
24 | 0.382 | 0.38 | -0.002 | -0.52356 | -0.52493 | 0 | 0.289 | -13.1488 | 18 | 0.381 | -0.52493 |  | 0.382 | 0.38 | 24 | 32160700
26 | 0.314 | 0.31 | -0.004 | -1.2739 | -1.2821 | 0 | 0.312 | -1.2821 | 19 | 0.312 | -1.2821 |  | 0.314 | 0.31 | 26 | 32160702
28 | 0.633 | 0.52 | -0.113 | -17.8515 | -19.601 | 0 | 0.3335 | -2.099 | 20 | 0.5765 | -19.601 |  | 0.633 | 0.52 | 28 | 38160723
31 | 1.69 | 1.73 | 0.04 | 2.3669 | 2.3392 | 0 | 0.3365 | -3.8633 | 21 | 1.71 | 2.3392 |  | 1.69 | 1.73 | 31 | 37254199
32 | 0.477 | 0.46 | -0.017 | -3.5639 | -3.6286 | 0 | 0.35 | -5.7143 | 22 | 0.4685 | -3.6286 |  | 0.477 | 0.46 | 32 | 37254857
33 | 0.921 | 0.51 | -0.411 | -44.6254 | -57.4423 | 0 | 0.376 | -24.4681 | 23 | 0.7155 | -57.4423 |  | 0.921 | 0.51 | 33 | 38162438
34 | 0.108 | 0.11 | 0.002 | 1.8519 | 1.8349 | 0 | 0.381 | -0.52493 | 24 | 0.109 | 1.8349 |  | 0.108 | 0.11 | 34 | 38162540
35 | 1.1 | 0.84 | -0.26 | -23.6364 | -26.8041 | 0 | 0.4525 | 3.3149 | 25 | 0.97 | -26.8041 |  | 1.1 | 0.84 | 35 | 38162433
36 | 0.218 | 0.19 | -0.028 | -12.844 | -13.7255 | 0 | 0.4655 | -15.2524 | 26 | 0.204 | -13.7255 |  | 0.218 | 0.19 | 36 | 37254297
37 | 0.003 | 0.05 | 0.047 | 1566.6667 | 177.3585 | 0 | 0.4685 | -3.6286 | 27 | 0.0265 | 177.3585 |  | 0.003 | 0.05 | 37 | 38163109
39 | 1.03 | 0.71 | -0.32 | -31.068 | -36.7816 | 0 | 0.557 | -27.6481 | 28 | 0.87 | -36.7816 |  | 1.03 | 0.71 | 39 | 38162503
42 | 0.03 | 0.06 | 0.03 | 100 | 66.6667 | 0 | 0.5765 | -19.601 | 29 | 0.045 | 66.6667 |  | 0.03 | 0.06 | 42 | 32136776
44 | 0.125 | 0.15 | 0.025 | 20 | 18.1818 | 0 | 0.61 | -6.5574 | 30 | 0.1375 | 18.1818 |  | 0.125 | 0.15 | 44 | 38165868
45 | 1.72 | 1.7 | -0.02 | -1.1628 | -1.1696 | 0 | 0.7155 | -57.4423 | 31 | 1.71 | -1.1696 |  | 1.72 | 1.7 | 45 | 37261409
46 | 0.343 | 0.33 | -0.013 | -3.7901 | -3.8633 | 0 | 0.7165 | -10.1884 | 32 | 0.3365 | -3.8633 |  | 0.343 | 0.33 | 46 | 37261865
47 | 1.02 | 1.03 | 0.01 | 0.98039 | 0.97561 | 0 | 0.787 | 0.76239 | 33 | 1.025 | 0.97561 |  | 1.02 | 1.03 | 47 | 37261874
48 | 0.009 | 0.04 | 0.031 | 344.4444 | 126.5306 | 0 | 0.87 | -36.7816 | 34 | 0.0245 | 126.5306 |  | 0.009 | 0.04 | 48 | 38165776
50 | 1.65 | 1.43 | -0.22 | -13.3333 | -14.2857 | 0 | 0.8955 | -19.0955 | 35 | 1.54 | -14.2857 |  | 1.65 | 1.43 | 50 | 38165773
51 | 0.003 | 0.04 | 0.037 | 1233.3333 | 172.093 | 0 | 0.94 | -19.1489 | 36 | 0.0215 | 172.093 |  | 0.003 | 0.04 | 51 | 38165845
52 | 1.03 | 0.85 | -0.18 | -17.4757 | -19.1489 | 0 | 0.97 | -26.8041 | 37 | 0.94 | -19.1489 |  | 1.03 | 0.85 | 52 | 32144696
53 | 0.337 | 0.33 | -0.007 | -2.0772 | -2.099 | 0 | 0.976 | 0.81967 | 38 | 0.3335 | -2.099 |  | 0.337 | 0.33 | 53 | 37264662
54 | 1.66 | 1.7 | 0.04 | 2.4096 | 2.381 | 0 | 1.025 | 0.97561 | 39 | 1.68 | 2.381 |  | 1.66 | 1.7 | 54 | 32023880
55 | 0.27 | 0.3 | 0.03 | 11.1111 | 10.5263 | 0 | 1.04 | 0 | 40 | 0.285 | 10.5263 |  | 0.27 | 0.3 | 55 | 37253832
56 | 0.003 | 0.02 | 0.017 | 566.6667 | 147.8261 | 0 | 1.045 | -8.6124 | 41 | 0.0115 | 147.8261 |  | 0.003 | 0.02 | 56 | 32023881
57 | 1.04 | 1.04 | 0 | 0 | 0 | 0 | 1.06 | -5.6604 | 42 | 1.04 | 0 |  | 1.04 | 1.04 | 57 | 38166192
58 | 0.784 | 0.79 | 0.006 | 0.76531 | 0.76239 | 0 | 1.065 | -2.8169 | 43 | 0.787 | 0.76239 |  | 0.784 | 0.79 | 58 | 32144698
59 | 0.634 | 0.48 | -0.154 | -24.2902 | -27.6481 | 0 | 1.28 | -18.75 | 44 | 0.557 | -27.6481 |  | 0.634 | 0.48 | 59 | 32023882
60 | 1.09 | 1.03 | -0.06 | -5.5046 | -5.6604 | 0 | 1.365 | 2.1978 | 45 | 1.06 | -5.6604 |  | 1.09 | 1.03 | 60 | 38166993
61 | 0.445 | 0.46 | 0.015 | 3.3708 | 3.3149 | 0 | 1.41 | 5.6738 | 46 | 0.4525 | 3.3149 |  | 0.445 | 0.46 | 61 | 32101002
66 | 0.02 | 0.06 | 0.04 | 200 | 100 | 0 | 1.455 | -17.1821 | 47 | 0.04 | 100 |  | 0.02 | 0.06 | 66 | 32144701
67 | 1.4 | 1.16 | -0.24 | -17.1429 | -18.75 | 0 | 1.54 | -14.2857 | 48 | 1.28 | -18.75 |  | 1.4 | 1.16 | 67 | 37272594
71 | 0.254 | 0.27 | 0.016 | 6.2992 | 6.1069 | 0 | 1.68 | 2.381 | 49 | 0.262 | 6.1069 |  | 0.254 | 0.27 | 71 | 37272564
78 | 1.35 | 1.38 | 0.03 | 2.2222 | 2.1978 | 0 | 1.71 | 2.3392 | 50 | 1.365 | 2.1978 |  | 1.35 | 1.38 | 78 | 32144703
89 | 1.37 | 1.45 | 0.08 | 5.8394 | 5.6738 | 0 | 1.71 | -1.1696 | 51 | 1.41 | 5.6738 |  | 1.37 | 1.45 | 89 | 32003034
Note on A114: ###res
Data Assessment
P/B Regression
Y = 0.920 * X + 0.020
md(95) = 0.133
N = 51, r = 0.981
Cusum test for linearity shows no significant deviation from linearity.
* indicates significant difference (rejection of null-hypothesis,
gamma = 0.05 for slope or intercept from P/B and for slope from SPC).

## Werte bis 10
Method X: PSA ng/ml |  |  | Instrument: | E 170
Method Y: PSA ng/ml |  |  | Instrument: | Centaur
82
Descriptive Statistics
| X | Y | Y-X | (Y - X)% of X
Median | 1.09 | 1.035 | -0.02 | -2.6363
Mean | 2.509 | 2.406 | -0.10304 | 44.2718
Minimum | 0.003 | 0.02 | -1.21 | -44.6254
Maximum | 8.72 | 8.76 | 0.61 | 1566.6667
68% Median Distance | 2.22 | 1.865 | 0.205 | 12.0437
Standard Deviation | 2.7091 | 2.6577 | 0.29238 | 230.8331
Differences
Medians |  |  |  | -5.0459
Means |  |  |  | -4.1067
Regression and Correlation Analysis
Coefficients of Correlation: | 0.99425 | 0.95002 |  |  |  |  | 1.09 | 1.035 | -0.02 | -2.6363 |  | 3.1
|  | slope b | intercept a | lower limit | upper limit |  | 2.509 | 2.406 | -0.10304 | 44.2718 |  | 2012-04-26 10:05
Structural Relationship Model: | [6 empty] | 0.003 | 0.02 | -1.21 | -44.6254 |  | PSA
Passing/Bablok (P/B) |  | 0.96011 | 0.0059095 |  |  |  | 8.72 | 8.76 | 0.61 | 1566.6667
95% Confidence Region for b (P/B) |  |  |  | 0.93716 | 0.98661 |  | 2.22 | 1.865 | 0.205 | 12.0437 |  | ng/ml
95% Confidence Region for a (P/B) |  |  |  | -0.010246 | 0.018605 |  | 2.7091 | 2.6577 | 0.29238 | 230.8331 |  | E 170
Std. Principal Component (SPC) |  | 0.98105 | -0.055483 |  |  |  | -5.0459 | -4.1067 | 0 | 0 |  | PSA
Linear Model: | [6 empty] | 0.99425 | 0.95002 | 0.99497 | 0
Least Squares Regression |  | 0.9754 | -0.041321 |  |  |  | 0.96011 | 1 | 0.93716 | 0.98661 |  | ng/ml
Theil Regression |  | 0.95536 | 0.0090297 |  |  |  | 0.0059095 | 0 | -0.010246 | 0.018605 |  | Centaur
Dispersion of Residuals: | [6 empty] | 0.98105 | 0 | -0.055483 | 0 |  | 82
Passing/Bablok Regression | md(68) = | 0.12629 | md(95) = | 0.45952 |  |  | 0.9754 | 0 | -0.041321 | 0
Std. Principal Component |  |  | SE = | 0.20482 |  |  | 1.0135 | 0 | 0.070664 | 0
[7 empty] | 0.95536 | 0 | 0.0090297 | 0 |  | 0.96011
[7 empty] | 0.12629 | 0.45952 | 0.20482 | 0 |  | Yes
[7 empty] | 0 | 0 | 1 | 82 |  | 0.0059095
0.12629
0.45952
0.99425
0.95002
1
2
1
2
TRUE
TRUE
0
10
0
10
1
1
1
1
1
1
FALSE
FALSE
0
90
18
TRUE
Statistics | Method X: PSA ng/ml | Method Y: PSA ng/ml
N | 82 | 82
Mean | 2.509 | 2.406
Median | 1.09 | 1.035
Minimum | 0.003 | 0.02
Maximum | 8.72 | 8.76
Range | 8.717 | 8.74
Difference Plot
(Normalized Differences)
| Sample Values | [18 empty] | 5 | 10 | [5 empty] | 5 | 20
Serial Number | X | Y | Y - X | (Y - X)% of X | Normalized Difference (%) | [14 empty] | 10 | 15 | [5 empty] | 10 | 25
[6 empty] | Diff Data Preset |  | Diff Ref Preset
scale_min | 0 | 0 | 0 |  |  | 0.4 | 0.05 | 0 | -0.05 |  |  |  | SelectData | [6 empty] | SelectCalibs | [6 empty] | SelectQCs | [6 empty] | F1_select |  | F2_select |  | grp_select
scale_max | 9 |  | 9 |  |  | 0.6 | -0.05 | 177.3585 | -0.05 |  |  |  | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? |  | #NAME?
| 0 |  |  | 0.0059095 |  |  |  | -57.4423 |  |  |  | data_count | 82 | [5 empty] | calib_count | 0 | [5 empty] | QC_count | 0 | [5 empty] | F1_count |  | F2_count |  | selFilt_count |  |  |  |  | selCrit |  |  | H | I | J | K | L | M
| 9 |  |  | 8.6469 | [8 empty] | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | filter1 |  | filter2 |  | Selected Filters | F1 | Grp Start | Legend |  | X_Vektor | Y_Vektor | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2
3 | 0.166 | 0.19 | 0.024 | 14.4578 | 13.4831 | 0 | 0.0115 | 147.8261 | 1 | 0.178 | 13.4831 |  | 0.166 | 0.19 | 3 | 38158035
4 | 1.08 | 1.05 | -0.03 | -2.7778 | -2.8169 | 0 | 0.0215 | 172.093 | 2 | 1.065 | -2.8169 |  | 1.08 | 1.05 | 4 | 38158883
5 | 0.63 | 0.59 | -0.04 | -6.3492 | -6.5574 | 0 | 0.0245 | 126.5306 | 3 | 0.61 | -6.5574 |  | 0.63 | 0.59 | 5 | 37239128
7 | 0.088 | 0.09 | 0.002 | 2.2727 | 2.2472 | 0 | 0.0265 | 177.3585 | 4 | 0.089 | 2.2472 |  | 0.088 | 0.09 | 7 | 37249463
8 | 0.36 | 0.34 | -0.02 | -5.5556 | -5.7143 | 0 | 0.04 | 100 | 5 | 0.35 | -5.7143 |  | 0.36 | 0.34 | 8 | 32127996
9 | 0.308 | 0.27 | -0.038 | -12.3377 | -13.1488 | 0 | 0.045 | 66.6667 | 6 | 0.289 | -13.1488 |  | 0.308 | 0.27 | 9 | 38159485
10 | 0.091 | 0.09 | -0.001 | -1.0989 | -1.105 | 0 | 0.072 | 22.2222 | 7 | 0.0905 | -1.105 |  | 0.091 | 0.09 | 10 | 32019594
11 | 1.09 | 1 | -0.09 | -8.2569 | -8.6124 | 0 | 0.089 | 2.2472 | 8 | 1.045 | -8.6124 |  | 1.09 | 1 | 11 | 37250625
12 | 0.972 | 0.98 | 0.008 | 0.82305 | 0.81967 | 0 | 0.0905 | -1.105 | 9 | 0.976 | 0.81967 |  | 0.972 | 0.98 | 12 | 38159496
13 | 0.753 | 0.68 | -0.073 | -9.6946 | -10.1884 | 0 | 0.109 | 1.8349 | 10 | 0.7165 | -10.1884 |  | 0.753 | 0.68 | 13 | 38158965
14 | 0.981 | 0.81 | -0.171 | -17.4312 | -19.0955 | 0 | 0.1375 | 18.1818 | 11 | 0.8955 | -19.0955 |  | 0.981 | 0.81 | 14 | 38159754
15 | 8.72 | 8.23 | -0.49 | -5.6193 | -5.7817 | 0 | 0.1535 | -17.5896 | 12 | 8.475 | -5.7817 |  | 8.72 | 8.23 | 15 | 32157083
16 | 0.167 | 0.14 | -0.027 | -16.1677 | -17.5896 | 0 | 0.178 | 13.4831 | 13 | 0.1535 | -17.5896 |  | 0.167 | 0.14 | 16 | 38158531
17 | 5 | 4.75 | -0.25 | -5 | -5.1282 | 0 | 0.204 | -13.7255 | 14 | 4.875 | -5.1282 |  | 5 | 4.75 | 17 | 38160092
18 | 0.501 | 0.43 | -0.071 | -14.1717 | -15.2524 | 0 | 0.2265 | -5.7395 | 15 | 0.4655 | -15.2524 |  | 0.501 | 0.43 | 18 | 37252133
19 | 0.422 | 0.33 | -0.092 | -21.8009 | -24.4681 | 0 | 0.262 | 6.1069 | 16 | 0.376 | -24.4681 |  | 0.422 | 0.33 | 19 | 32023925
20 | 0.233 | 0.22 | -0.013 | -5.5794 | -5.7395 | 0 | 0.285 | 10.5263 | 17 | 0.2265 | -5.7395 |  | 0.233 | 0.22 | 20 | 38160098
21 | 5.22 | 4.2 | -1.02 | -19.5402 | -21.6561 | 0 | 0.289 | -13.1488 | 18 | 4.71 | -21.6561 |  | 5.22 | 4.2 | 21 | 32121489
22 | 1.58 | 1.33 | -0.25 | -15.8228 | -17.1821 | 0 | 0.312 | -1.2821 | 19 | 1.455 | -17.1821 |  | 1.58 | 1.33 | 22 | 37252141
23 | 0.064 | 0.08 | 0.016 | 25 | 22.2222 | 0 | 0.3335 | -2.099 | 20 | 0.072 | 22.2222 |  | 0.064 | 0.08 | 23 | 37252658
24 | 0.382 | 0.38 | -0.002 | -0.52356 | -0.52493 | 0 | 0.3365 | -3.8633 | 21 | 0.381 | -0.52493 |  | 0.382 | 0.38 | 24 | 32160700
26 | 0.314 | 0.31 | -0.004 | -1.2739 | -1.2821 | 0 | 0.35 | -5.7143 | 22 | 0.312 | -1.2821 |  | 0.314 | 0.31 | 26 | 32160702
28 | 0.633 | 0.52 | -0.113 | -17.8515 | -19.601 | 0 | 0.376 | -24.4681 | 23 | 0.5765 | -19.601 |  | 0.633 | 0.52 | 28 | 38160723
31 | 1.69 | 1.73 | 0.04 | 2.3669 | 2.3392 | 0 | 0.381 | -0.52493 | 24 | 1.71 | 2.3392 |  | 1.69 | 1.73 | 31 | 37254199
32 | 0.477 | 0.46 | -0.017 | -3.5639 | -3.6286 | 0 | 0.4525 | 3.3149 | 25 | 0.4685 | -3.6286 |  | 0.477 | 0.46 | 32 | 37254857
33 | 0.921 | 0.51 | -0.411 | -44.6254 | -57.4423 | 0 | 0.4655 | -15.2524 | 26 | 0.7155 | -57.4423 |  | 0.921 | 0.51 | 33 | 38162438
34 | 0.108 | 0.11 | 0.002 | 1.8519 | 1.8349 | 0 | 0.4685 | -3.6286 | 27 | 0.109 | 1.8349 |  | 0.108 | 0.11 | 34 | 38162540
35 | 1.1 | 0.84 | -0.26 | -23.6364 | -26.8041 | 0 | 0.557 | -27.6481 | 28 | 0.97 | -26.8041 |  | 1.1 | 0.84 | 35 | 38162433
36 | 0.218 | 0.19 | -0.028 | -12.844 | -13.7255 | 0 | 0.5765 | -19.601 | 29 | 0.204 | -13.7255 |  | 0.218 | 0.19 | 36 | 37254297
37 | 0.003 | 0.05 | 0.047 | 1566.6667 | 177.3585 | 0 | 0.61 | -6.5574 | 30 | 0.0265 | 177.3585 |  | 0.003 | 0.05 | 37 | 38163109
38 | 8.38 | 8.03 | -0.35 | -4.1766 | -4.2657 | 0 | 0.7155 | -57.4423 | 31 | 8.205 | -4.2657 |  | 8.38 | 8.03 | 38 | 38162501
39 | 1.03 | 0.71 | -0.32 | -31.068 | -36.7816 | 0 | 0.7165 | -10.1884 | 32 | 0.87 | -36.7816 |  | 1.03 | 0.71 | 39 | 38162503
40 | 4.81 | 4.69 | -0.12 | -2.4948 | -2.5263 | 0 | 0.787 | 0.76239 | 33 | 4.75 | -2.5263 |  | 4.81 | 4.69 | 40 | 37264171
41 | 5.21 | 4 | -1.21 | -23.2246 | -26.2758 | 0 | 0.87 | -36.7816 | 34 | 4.605 | -26.2758 |  | 5.21 | 4 | 41 | 32160738
42 | 0.03 | 0.06 | 0.03 | 100 | 66.6667 | 0 | 0.8955 | -19.0955 | 35 | 0.045 | 66.6667 |  | 0.03 | 0.06 | 42 | 32136776
43 | 8.25 | 7.67 | -0.58 | -7.0303 | -7.2864 | 0 | 0.94 | -19.1489 | 36 | 7.96 | -7.2864 |  | 8.25 | 7.67 | 43 | 32127959
44 | 0.125 | 0.15 | 0.025 | 20 | 18.1818 | 0 | 0.97 | -26.8041 | 37 | 0.1375 | 18.1818 |  | 0.125 | 0.15 | 44 | 38165868
45 | 1.72 | 1.7 | -0.02 | -1.1628 | -1.1696 | 0 | 0.976 | 0.81967 | 38 | 1.71 | -1.1696 |  | 1.72 | 1.7 | 45 | 37261409
46 | 0.343 | 0.33 | -0.013 | -3.7901 | -3.8633 | 0 | 1.025 | 0.97561 | 39 | 0.3365 | -3.8633 |  | 0.343 | 0.33 | 46 | 37261865
47 | 1.02 | 1.03 | 0.01 | 0.98039 | 0.97561 | 0 | 1.04 | 0 | 40 | 1.025 | 0.97561 |  | 1.02 | 1.03 | 47 | 37261874
48 | 0.009 | 0.04 | 0.031 | 344.4444 | 126.5306 | 0 | 1.045 | -8.6124 | 41 | 0.0245 | 126.5306 |  | 0.009 | 0.04 | 48 | 38165776
50 | 1.65 | 1.43 | -0.22 | -13.3333 | -14.2857 | 0 | 1.06 | -5.6604 | 42 | 1.54 | -14.2857 |  | 1.65 | 1.43 | 50 | 38165773
51 | 0.003 | 0.04 | 0.037 | 1233.3333 | 172.093 | 0 | 1.065 | -2.8169 | 43 | 0.0215 | 172.093 |  | 0.003 | 0.04 | 51 | 38165845
52 | 1.03 | 0.85 | -0.18 | -17.4757 | -19.1489 | 0 | 1.28 | -18.75 | 44 | 0.94 | -19.1489 |  | 1.03 | 0.85 | 52 | 32144696
53 | 0.337 | 0.33 | -0.007 | -2.0772 | -2.099 | 0 | 1.365 | 2.1978 | 45 | 0.3335 | -2.099 |  | 0.337 | 0.33 | 53 | 37264662
54 | 1.66 | 1.7 | 0.04 | 2.4096 | 2.381 | 0 | 1.41 | 5.6738 | 46 | 1.68 | 2.381 |  | 1.66 | 1.7 | 54 | 32023880
55 | 0.27 | 0.3 | 0.03 | 11.1111 | 10.5263 | 0 | 1.455 | -17.1821 | 47 | 0.285 | 10.5263 |  | 0.27 | 0.3 | 55 | 37253832
56 | 0.003 | 0.02 | 0.017 | 566.6667 | 147.8261 | 0 | 1.54 | -14.2857 | 48 | 0.0115 | 147.8261 |  | 0.003 | 0.02 | 56 | 32023881
57 | 1.04 | 1.04 | 0 | 0 | 0 | 0 | 1.68 | 2.381 | 49 | 1.04 | 0 |  | 1.04 | 1.04 | 57 | 38166192
58 | 0.784 | 0.79 | 0.006 | 0.76531 | 0.76239 | 0 | 1.71 | 2.3392 | 50 | 0.787 | 0.76239 |  | 0.784 | 0.79 | 58 | 32144698
59 | 0.634 | 0.48 | -0.154 | -24.2902 | -27.6481 | 0 | 1.71 | -1.1696 | 51 | 0.557 | -27.6481 |  | 0.634 | 0.48 | 59 | 32023882
60 | 1.09 | 1.03 | -0.06 | -5.5046 | -5.6604 | 0 | 2.01 | -2.9851 | 52 | 1.06 | -5.6604 |  | 1.09 | 1.03 | 60 | 38166993
61 | 0.445 | 0.46 | 0.015 | 3.3708 | 3.3149 | 0 | 2.28 | -7.0175 | 53 | 0.4525 | 3.3149 |  | 0.445 | 0.46 | 61 | 32101002
63 | 5.28 | 4.51 | -0.77 | -14.5833 | -15.7303 | 0 | 2.635 | 4.1746 | 54 | 4.895 | -15.7303 |  | 5.28 | 4.51 | 63 | 32160744
64 | 5.26 | 5.04 | -0.22 | -4.1825 | -4.2718 | 0 | 2.72 | -21.3235 | 55 | 5.15 | -4.2718 |  | 5.26 | 5.04 | 64 | 38167182
66 | 0.02 | 0.06 | 0.04 | 200 | 100 | 0 | 2.73 | 0 | 56 | 0.04 | 100 |  | 0.02 | 0.06 | 66 | 32144701
67 | 1.4 | 1.16 | -0.24 | -17.1429 | -18.75 | 0 | 3.34 | -16.1677 | 57 | 1.28 | -18.75 |  | 1.4 | 1.16 | 67 | 37272594
69 | 8.09 | 7.46 | -0.63 | -7.7874 | -8.1029 | 0 | 3.705 | -6.7476 | 58 | 7.775 | -8.1029 |  | 8.09 | 7.46 | 69 | 32160613
70 | 4.84 | 5.3 | 0.46 | 9.5041 | 9.073 | 0 | 4.4 | -1.8182 | 59 | 5.07 | 9.073 |  | 4.84 | 5.3 | 70 | 38170528
71 | 0.254 | 0.27 | 0.016 | 6.2992 | 6.1069 | 0 | 4.605 | -26.2758 | 60 | 0.262 | 6.1069 |  | 0.254 | 0.27 | 71 | 37272564
72 | 7.59 | 7.92 | 0.33 | 4.3478 | 4.2553 | 0 | 4.71 | -21.6561 | 61 | 7.755 | 4.2553 |  | 7.59 | 7.92 | 72 | 38170067
76 | 5.21 | 5.53 | 0.32 | 6.142 | 5.959 | 0 | 4.75 | -2.5263 | 62 | 5.37 | 5.959 |  | 5.21 | 5.53 | 76 | 38171774
78 | 1.35 | 1.38 | 0.03 | 2.2222 | 2.1978 | 0 | 4.875 | -5.1282 | 63 | 1.365 | 2.1978 |  | 1.35 | 1.38 | 78 | 32144703
79 | 8.15 | 8.76 | 0.61 | 7.4847 | 7.2147 | 0 | 4.895 | -15.7303 | 64 | 8.455 | 7.2147 |  | 8.15 | 8.76 | 79 | 32160631
80 | 7.36 | 7.63 | 0.27 | 3.6685 | 3.6024 | 0 | 4.965 | -10.6747 | 65 | 7.495 | 3.6024 |  | 7.36 | 7.63 | 80 | 32160629
85 | 5.23 | 4.7 | -0.53 | -10.1338 | -10.6747 | 0 | 5.07 | 9.073 | 66 | 4.965 | -10.6747 |  | 5.23 | 4.7 | 85 | 32160633
86 | 6.75 | 6.69 | -0.06 | -0.88889 | -0.89286 | 0 | 5.15 | -4.2718 | 67 | 6.72 | -0.89286 |  | 6.75 | 6.69 | 86 | 37284567
89 | 1.37 | 1.45 | 0.08 | 5.8394 | 5.6738 | 0 | 5.37 | 5.959 | 68 | 1.41 | 5.6738 |  | 1.37 | 1.45 | 89 | 32003034
90 | 7.08 | 7.25 | 0.17 | 2.4011 | 2.3726 | 0 | 5.48 | 1.0949 | 69 | 7.165 | 2.3726 |  | 7.08 | 7.25 | 90 | 37287086
91 | 7.93 | 7.42 | -0.51 | -6.4313 | -6.645 | 0 | 5.935 | -10.278 | 70 | 7.675 | -6.645 |  | 7.93 | 7.42 | 91 | 38176759
92 | 2.73 | 2.73 | 0 | 0 | 0 | 0 | 6.125 | 3.102 | 71 | 2.73 | 0 |  | 2.73 | 2.73 | 92
93 | 2.04 | 1.98 | -0.06 | -2.9412 | -2.9851 | 0 | 6.345 | 7.4074 | 72 | 2.01 | -2.9851 |  | 2.04 | 1.98 | 93
94 | 4.44 | 4.36 | -0.08 | -1.8018 | -1.8182 | 0 | 6.72 | -0.89286 | 73 | 4.4 | -1.8182 |  | 4.44 | 4.36 | 94
95 | 6.24 | 5.63 | -0.61 | -9.7756 | -10.278 | 0 | 7.165 | 2.3726 | 74 | 5.935 | -10.278 |  | 6.24 | 5.63 | 95
96 | 3.83 | 3.58 | -0.25 | -6.5274 | -6.7476 | 0 | 7.495 | 3.6024 | 75 | 3.705 | -6.7476 |  | 3.83 | 3.58 | 96
97 | 3.01 | 2.43 | -0.58 | -19.2691 | -21.3235 | 0 | 7.675 | -6.645 | 76 | 2.72 | -21.3235 |  | 3.01 | 2.43 | 97
98 | 2.36 | 2.2 | -0.16 | -6.7797 | -7.0175 | 0 | 7.755 | 4.2553 | 77 | 2.28 | -7.0175 |  | 2.36 | 2.2 | 98
99 | 3.61 | 3.07 | -0.54 | -14.9584 | -16.1677 | 0 | 7.775 | -8.1029 | 78 | 3.34 | -16.1677 |  | 3.61 | 3.07 | 99
100 | 6.11 | 6.58 | 0.47 | 7.6923 | 7.4074 | 0 | 7.96 | -7.2864 | 79 | 6.345 | 7.4074 |  | 6.11 | 6.58 | 100
101 | 6.03 | 6.22 | 0.19 | 3.1509 | 3.102 | 0 | 8.205 | -4.2657 | 80 | 6.125 | 3.102 |  | 6.03 | 6.22 | 101
102 | 5.45 | 5.51 | 0.06 | 1.1009 | 1.0949 | 0 | 8.455 | 7.2147 | 81 | 5.48 | 1.0949 |  | 5.45 | 5.51 | 102
103 | 2.58 | 2.69 | 0.11 | 4.2636 | 4.1746 | 0 | 8.475 | -5.7817 | 82 | 2.635 | 4.1746 |  | 2.58 | 2.69 | 103
Note on A114: ###res
Data Assessment
P/B Regression
Y = 0.960 * X + 0.006
md(95) = 0.460
N = 82, r = 0.994
Cusum test for linearity shows no significant deviation from linearity.
* indicates significant difference (rejection of null-hypothesis,
gamma = 0.05 for slope or intercept from P/B and for slope from SPC).

## alle Werte
Method X: PSA ng/ml |  |  | Instrument: | E 170
Method Y: PSA ng/ml |  |  | Instrument: | Centaur
102
Descriptive Statistics
| X | Y | Y-X | (Y - X)% of X
Median | 1.88 | 1.855 | -0.05 | -3.9833
Mean | 9.6494 | 8.9433 | -0.70607 | 34.1176
Minimum | 0.003 | 0.02 | -15.4 | -44.6254
Maximum | 81.4 | 75.25 | 1.95 | 1566.6667
68% Median Distance | 3.86 | 3.725 | 0.45 | 10.4413
Standard Deviation | 16.8832 | 15.4357 | 2.1759 | 207.7796
Differences
Medians |  |  |  | -1.3298
Means |  |  |  | -7.3172
Regression and Correlation Analysis
Coefficients of Correlation: | 0.99494 | 0.96311 |  |  |  |  | 1.88 | 1.855 | -0.05 | -3.9833 |  | 3.1
|  | slope b | intercept a | lower limit | upper limit |  | 9.6494 | 8.9433 | -0.70607 | 34.1176 |  | 2012-04-26 10:04
Structural Relationship Model: | [6 empty] | 0.003 | 0.02 | -15.4 | -44.6254 |  | PSA
Passing/Bablok (P/B) |  | 0.94107 | 0.013682 |  |  |  | 81.4 | 75.25 | 1.95 | 1566.6667
95% Confidence Region for b (P/B) |  |  |  | 0.92577 | 0.96022 |  | 3.86 | 3.725 | 0.45 | 10.4413 |  | ng/ml
95% Confidence Region for a (P/B) |  |  |  | -0.0015432 | 0.031947 |  | 16.8832 | 15.4357 | 2.1759 | 207.7796 |  | E 170
Std. Principal Component (SPC) |  | 0.91427 | 0.12122 |  |  |  | -1.3298 | -7.3172 | 0 | 0 |  | PSA
Linear Model: | [6 empty] | 0.99494 | 0.96311 | 0.99716 | 0
Least Squares Regression |  | 0.90964 | 0.1659 |  |  |  | 0.94107 | 1 | 0.92577 | 0.96022 |  | ng/ml
Theil Regression |  | 0.93839 | 0.014555 |  |  |  | 0.013682 | 0 | -0.0015432 | 0.031947 |  | Centaur
Dispersion of Residuals: | [6 empty] | 0.91427 | 1 | 0.12122 | 0 |  | 102
Passing/Bablok Regression | md(68) = | 0.20825 | md(95) = | 2.3124 |  |  | 0.90964 | 0 | 0.1659 | 0
Std. Principal Component |  |  | SE = | 1.1568 |  |  | 1.0882 | 0 | -0.083055 | 0
[7 empty] | 0.93839 | 0 | 0.014555 | 0 |  | 0.94107
[7 empty] | 0.20825 | 2.3124 | 1.1568 | 0 |  | Yes
[7 empty] | 0 | 0 | 1 | 102 |  | 0.013682
0.20825
2.3124
0.99494
0.96311
1
2
1
2
TRUE
TRUE
0
0
1
1
1
1
1
1
FALSE
FALSE
0
9
1.8
TRUE
Statistics | Method X: PSA ng/ml | Method Y: PSA ng/ml
N | 102 | 102
Mean | 9.6494 | 8.9433
Median | 1.88 | 1.855
Minimum | 0.003 | 0.02
Maximum | 81.4 | 75.25
Range | 81.397 | 75.23
Difference Plot
(Normalized Differences)
| Sample Values | [18 empty] | 5 | 10 | [5 empty] | 5 | 20
Serial Number | X | Y | Y - X | (Y - X)% of X | Normalized Difference (%) | [14 empty] | 10 | 15 | [5 empty] | 10 | 25
[6 empty] | Diff Data Preset |  | Diff Ref Preset
scale_min | 0 | 0 | 0 |  |  | 0.4 | 0.05 | 0 | -0.05 |  |  |  | SelectData | [6 empty] | SelectCalibs | [6 empty] | SelectQCs | [6 empty] | F1_select |  | F2_select |  | grp_select
scale_max | 90 |  | 90 |  |  | 0.6 | -0.05 | 177.3585 | -0.05 |  |  |  | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? |  | #NAME?
| 0 |  |  | 0.013682 |  |  |  | -57.4423 |  |  |  | data_count | 102 | [5 empty] | calib_count | 0 | [5 empty] | QC_count | 0 | [5 empty] | F1_count |  | F2_count |  | selFilt_count |  |  |  |  | selCrit |  |  | H | I | J | K | L | M
| 90 |  |  | 84.7098 | [8 empty] | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | filter1 |  | filter2 |  | Selected Filters | F1 | Grp Start | Legend |  | X_Vektor | Y_Vektor | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2
1 | 18 | 15.54 | -2.46 | -13.6667 | -14.6691 | 0 | 0.0115 | 147.8261 | 1 | 16.77 | -14.6691 |  | 18 | 15.54 | 1 | 38158289
2 | 27.4 | 25.5 | -1.9 | -6.9343 | -7.1834 | 0 | 0.0215 | 172.093 | 2 | 26.45 | -7.1834 |  | 27.4 | 25.5 | 2 | 38158688
3 | 0.166 | 0.19 | 0.024 | 14.4578 | 13.4831 | 0 | 0.0245 | 126.5306 | 3 | 0.178 | 13.4831 |  | 0.166 | 0.19 | 3 | 38158035
4 | 1.08 | 1.05 | -0.03 | -2.7778 | -2.8169 | 0 | 0.0265 | 177.3585 | 4 | 1.065 | -2.8169 |  | 1.08 | 1.05 | 4 | 38158883
5 | 0.63 | 0.59 | -0.04 | -6.3492 | -6.5574 | 0 | 0.04 | 100 | 5 | 0.61 | -6.5574 |  | 0.63 | 0.59 | 5 | 37239128
6 | 49.9 | 44.53 | -5.37 | -10.7615 | -11.3735 | 0 | 0.045 | 66.6667 | 6 | 47.215 | -11.3735 |  | 49.9 | 44.53 | 6 | 38158959
7 | 0.088 | 0.09 | 0.002 | 2.2727 | 2.2472 | 0 | 0.072 | 22.2222 | 7 | 0.089 | 2.2472 |  | 0.088 | 0.09 | 7 | 37249463
8 | 0.36 | 0.34 | -0.02 | -5.5556 | -5.7143 | 0 | 0.089 | 2.2472 | 8 | 0.35 | -5.7143 |  | 0.36 | 0.34 | 8 | 32127996
9 | 0.308 | 0.27 | -0.038 | -12.3377 | -13.1488 | 0 | 0.0905 | -1.105 | 9 | 0.289 | -13.1488 |  | 0.308 | 0.27 | 9 | 38159485
10 | 0.091 | 0.09 | -0.001 | -1.0989 | -1.105 | 0 | 0.109 | 1.8349 | 10 | 0.0905 | -1.105 |  | 0.091 | 0.09 | 10 | 32019594
11 | 1.09 | 1 | -0.09 | -8.2569 | -8.6124 | 0 | 0.1375 | 18.1818 | 11 | 1.045 | -8.6124 |  | 1.09 | 1 | 11 | 37250625
12 | 0.972 | 0.98 | 0.008 | 0.82305 | 0.81967 | 0 | 0.1535 | -17.5896 | 12 | 0.976 | 0.81967 |  | 0.972 | 0.98 | 12 | 38159496
13 | 0.753 | 0.68 | -0.073 | -9.6946 | -10.1884 | 0 | 0.178 | 13.4831 | 13 | 0.7165 | -10.1884 |  | 0.753 | 0.68 | 13 | 38158965
14 | 0.981 | 0.81 | -0.171 | -17.4312 | -19.0955 | 0 | 0.204 | -13.7255 | 14 | 0.8955 | -19.0955 |  | 0.981 | 0.81 | 14 | 38159754
15 | 8.72 | 8.23 | -0.49 | -5.6193 | -5.7817 | 0 | 0.2265 | -5.7395 | 15 | 8.475 | -5.7817 |  | 8.72 | 8.23 | 15 | 32157083
16 | 0.167 | 0.14 | -0.027 | -16.1677 | -17.5896 | 0 | 0.262 | 6.1069 | 16 | 0.1535 | -17.5896 |  | 0.167 | 0.14 | 16 | 38158531
17 | 5 | 4.75 | -0.25 | -5 | -5.1282 | 0 | 0.285 | 10.5263 | 17 | 4.875 | -5.1282 |  | 5 | 4.75 | 17 | 38160092
18 | 0.501 | 0.43 | -0.071 | -14.1717 | -15.2524 | 0 | 0.289 | -13.1488 | 18 | 0.4655 | -15.2524 |  | 0.501 | 0.43 | 18 | 37252133
19 | 0.422 | 0.33 | -0.092 | -21.8009 | -24.4681 | 0 | 0.312 | -1.2821 | 19 | 0.376 | -24.4681 |  | 0.422 | 0.33 | 19 | 32023925
20 | 0.233 | 0.22 | -0.013 | -5.5794 | -5.7395 | 0 | 0.3335 | -2.099 | 20 | 0.2265 | -5.7395 |  | 0.233 | 0.22 | 20 | 38160098
21 | 5.22 | 4.2 | -1.02 | -19.5402 | -21.6561 | 0 | 0.3365 | -3.8633 | 21 | 4.71 | -21.6561 |  | 5.22 | 4.2 | 21 | 32121489
22 | 1.58 | 1.33 | -0.25 | -15.8228 | -17.1821 | 0 | 0.35 | -5.7143 | 22 | 1.455 | -17.1821 |  | 1.58 | 1.33 | 22 | 37252141
23 | 0.064 | 0.08 | 0.016 | 25 | 22.2222 | 0 | 0.376 | -24.4681 | 23 | 0.072 | 22.2222 |  | 0.064 | 0.08 | 23 | 37252658
24 | 0.382 | 0.38 | -0.002 | -0.52356 | -0.52493 | 0 | 0.381 | -0.52493 | 24 | 0.381 | -0.52493 |  | 0.382 | 0.38 | 24 | 32160700
25 | 50.1 | 44.52 | -5.58 | -11.1377 | -11.7945 | 0 | 0.4525 | 3.3149 | 25 | 47.31 | -11.7945 |  | 50.1 | 44.52 | 25 | 32157144
26 | 0.314 | 0.31 | -0.004 | -1.2739 | -1.2821 | 0 | 0.4655 | -15.2524 | 26 | 0.312 | -1.2821 |  | 0.314 | 0.31 | 26 | 32160702
27 | 17.2 | 16.07 | -1.13 | -6.5698 | -6.7929 | 0 | 0.4685 | -3.6286 | 27 | 16.635 | -6.7929 |  | 17.2 | 16.07 | 27 | 32160703
28 | 0.633 | 0.52 | -0.113 | -17.8515 | -19.601 | 0 | 0.557 | -27.6481 | 28 | 0.5765 | -19.601 |  | 0.633 | 0.52 | 28 | 38160723
29 | 65.9 | 50.5 | -15.4 | -23.3687 | -26.4605 | 0 | 0.5765 | -19.601 | 29 | 58.2 | -26.4605 |  | 65.9 | 50.5 | 29 | 32160704
30 | 41 | 32.48 | -8.52 | -20.7805 | -23.19 | 0 | 0.61 | -6.5574 | 30 | 36.74 | -23.19 |  | 41 | 32.48 | 30 | 32160724
31 | 1.69 | 1.73 | 0.04 | 2.3669 | 2.3392 | 0 | 0.7155 | -57.4423 | 31 | 1.71 | 2.3392 |  | 1.69 | 1.73 | 31 | 37254199
32 | 0.477 | 0.46 | -0.017 | -3.5639 | -3.6286 | 0 | 0.7165 | -10.1884 | 32 | 0.4685 | -3.6286 |  | 0.477 | 0.46 | 32 | 37254857
33 | 0.921 | 0.51 | -0.411 | -44.6254 | -57.4423 | 0 | 0.787 | 0.76239 | 33 | 0.7155 | -57.4423 |  | 0.921 | 0.51 | 33 | 38162438
34 | 0.108 | 0.11 | 0.002 | 1.8519 | 1.8349 | 0 | 0.87 | -36.7816 | 34 | 0.109 | 1.8349 |  | 0.108 | 0.11 | 34 | 38162540
35 | 1.1 | 0.84 | -0.26 | -23.6364 | -26.8041 | 0 | 0.8955 | -19.0955 | 35 | 0.97 | -26.8041 |  | 1.1 | 0.84 | 35 | 38162433
36 | 0.218 | 0.19 | -0.028 | -12.844 | -13.7255 | 0 | 0.94 | -19.1489 | 36 | 0.204 | -13.7255 |  | 0.218 | 0.19 | 36 | 37254297
37 | 0.003 | 0.05 | 0.047 | 1566.6667 | 177.3585 | 0 | 0.97 | -26.8041 | 37 | 0.0265 | 177.3585 |  | 0.003 | 0.05 | 37 | 38163109
38 | 8.38 | 8.03 | -0.35 | -4.1766 | -4.2657 | 0 | 0.976 | 0.81967 | 38 | 8.205 | -4.2657 |  | 8.38 | 8.03 | 38 | 38162501
39 | 1.03 | 0.71 | -0.32 | -31.068 | -36.7816 | 0 | 1.025 | 0.97561 | 39 | 0.87 | -36.7816 |  | 1.03 | 0.71 | 39 | 38162503
40 | 4.81 | 4.69 | -0.12 | -2.4948 | -2.5263 | 0 | 1.04 | 0 | 40 | 4.75 | -2.5263 |  | 4.81 | 4.69 | 40 | 37264171
41 | 5.21 | 4 | -1.21 | -23.2246 | -26.2758 | 0 | 1.045 | -8.6124 | 41 | 4.605 | -26.2758 |  | 5.21 | 4 | 41 | 32160738
42 | 0.03 | 0.06 | 0.03 | 100 | 66.6667 | 0 | 1.06 | -5.6604 | 42 | 0.045 | 66.6667 |  | 0.03 | 0.06 | 42 | 32136776
43 | 8.25 | 7.67 | -0.58 | -7.0303 | -7.2864 | 0 | 1.065 | -2.8169 | 43 | 7.96 | -7.2864 |  | 8.25 | 7.67 | 43 | 32127959
44 | 0.125 | 0.15 | 0.025 | 20 | 18.1818 | 0 | 1.28 | -18.75 | 44 | 0.1375 | 18.1818 |  | 0.125 | 0.15 | 44 | 38165868
45 | 1.72 | 1.7 | -0.02 | -1.1628 | -1.1696 | 0 | 1.365 | 2.1978 | 45 | 1.71 | -1.1696 |  | 1.72 | 1.7 | 45 | 37261409
46 | 0.343 | 0.33 | -0.013 | -3.7901 | -3.8633 | 0 | 1.41 | 5.6738 | 46 | 0.3365 | -3.8633 |  | 0.343 | 0.33 | 46 | 37261865
47 | 1.02 | 1.03 | 0.01 | 0.98039 | 0.97561 | 0 | 1.455 | -17.1821 | 47 | 1.025 | 0.97561 |  | 1.02 | 1.03 | 47 | 37261874
48 | 0.009 | 0.04 | 0.031 | 344.4444 | 126.5306 | 0 | 1.54 | -14.2857 | 48 | 0.0245 | 126.5306 |  | 0.009 | 0.04 | 48 | 38165776
49 | 40.2 | 39.02 | -1.18 | -2.9353 | -2.979 | 0 | 1.68 | 2.381 | 49 | 39.61 | -2.979 |  | 40.2 | 39.02 | 49 | 38165229
50 | 1.65 | 1.43 | -0.22 | -13.3333 | -14.2857 | 0 | 1.71 | 2.3392 | 50 | 1.54 | -14.2857 |  | 1.65 | 1.43 | 50 | 38165773
51 | 0.003 | 0.04 | 0.037 | 1233.3333 | 172.093 | 0 | 1.71 | -1.1696 | 51 | 0.0215 | 172.093 |  | 0.003 | 0.04 | 51 | 38165845
52 | 1.03 | 0.85 | -0.18 | -17.4757 | -19.1489 | 0 | 2.01 | -2.9851 | 52 | 0.94 | -19.1489 |  | 1.03 | 0.85 | 52 | 32144696
53 | 0.337 | 0.33 | -0.007 | -2.0772 | -2.099 | 0 | 2.28 | -7.0175 | 53 | 0.3335 | -2.099 |  | 0.337 | 0.33 | 53 | 37264662
54 | 1.66 | 1.7 | 0.04 | 2.4096 | 2.381 | 0 | 2.635 | 4.1746 | 54 | 1.68 | 2.381 |  | 1.66 | 1.7 | 54 | 32023880
55 | 0.27 | 0.3 | 0.03 | 11.1111 | 10.5263 | 0 | 2.72 | -21.3235 | 55 | 0.285 | 10.5263 |  | 0.27 | 0.3 | 55 | 37253832
56 | 0.003 | 0.02 | 0.017 | 566.6667 | 147.8261 | 0 | 2.73 | 0 | 56 | 0.0115 | 147.8261 |  | 0.003 | 0.02 | 56 | 32023881
57 | 1.04 | 1.04 | 0 | 0 | 0 | 0 | 3.34 | -16.1677 | 57 | 1.04 | 0 |  | 1.04 | 1.04 | 57 | 38166192
58 | 0.784 | 0.79 | 0.006 | 0.76531 | 0.76239 | 0 | 3.705 | -6.7476 | 58 | 0.787 | 0.76239 |  | 0.784 | 0.79 | 58 | 32144698
59 | 0.634 | 0.48 | -0.154 | -24.2902 | -27.6481 | 0 | 4.4 | -1.8182 | 59 | 0.557 | -27.6481 |  | 0.634 | 0.48 | 59 | 32023882
60 | 1.09 | 1.03 | -0.06 | -5.5046 | -5.6604 | 0 | 4.605 | -26.2758 | 60 | 1.06 | -5.6604 |  | 1.09 | 1.03 | 60 | 38166993
61 | 0.445 | 0.46 | 0.015 | 3.3708 | 3.3149 | 0 | 4.71 | -21.6561 | 61 | 0.4525 | 3.3149 |  | 0.445 | 0.46 | 61 | 32101002
62 | 22.8 | 21.48 | -1.32 | -5.7895 | -5.9621 | 0 | 4.75 | -2.5263 | 62 | 22.14 | -5.9621 |  | 22.8 | 21.48 | 62 | 32137484
63 | 5.28 | 4.51 | -0.77 | -14.5833 | -15.7303 | 0 | 4.875 | -5.1282 | 63 | 4.895 | -15.7303 |  | 5.28 | 4.51 | 63 | 32160744
64 | 5.26 | 5.04 | -0.22 | -4.1825 | -4.2718 | 0 | 4.895 | -15.7303 | 64 | 5.15 | -4.2718 |  | 5.26 | 5.04 | 64 | 38167182
65 | 46.8 | 43.26 | -3.54 | -7.5641 | -7.8614 | 0 | 4.965 | -10.6747 | 65 | 45.03 | -7.8614 |  | 46.8 | 43.26 | 65 | 32144700
66 | 0.02 | 0.06 | 0.04 | 200 | 100 | 0 | 5.07 | 9.073 | 66 | 0.04 | 100 |  | 0.02 | 0.06 | 66 | 32144701
67 | 1.4 | 1.16 | -0.24 | -17.1429 | -18.75 | 0 | 5.15 | -4.2718 | 67 | 1.28 | -18.75 |  | 1.4 | 1.16 | 67 | 37272594
68 | 19.1 | 20.66 | 1.56 | 8.1675 | 7.8471 | 0 | 5.37 | 5.959 | 68 | 19.88 | 7.8471 |  | 19.1 | 20.66 | 68 | 38169363
69 | 8.09 | 7.46 | -0.63 | -7.7874 | -8.1029 | 0 | 5.48 | 1.0949 | 69 | 7.775 | -8.1029 |  | 8.09 | 7.46 | 69 | 32160613
70 | 4.84 | 5.3 | 0.46 | 9.5041 | 9.073 | 0 | 5.935 | -10.278 | 70 | 5.07 | 9.073 |  | 4.84 | 5.3 | 70 | 38170528
71 | 0.254 | 0.27 | 0.016 | 6.2992 | 6.1069 | 0 | 6.125 | 3.102 | 71 | 0.262 | 6.1069 |  | 0.254 | 0.27 | 71 | 37272564
72 | 7.59 | 7.92 | 0.33 | 4.3478 | 4.2553 | 0 | 6.345 | 7.4074 | 72 | 7.755 | 4.2553 |  | 7.59 | 7.92 | 72 | 38170067
73 | 53.9 | 54.88 | 0.98 | 1.8182 | 1.8018 | 0 | 6.72 | -0.89286 | 73 | 54.39 | 1.8018 |  | 53.9 | 54.88 | 73 | 32160615
74 | 56.7 | 57.33 | 0.63 | 1.1111 | 1.105 | 0 | 7.165 | 2.3726 | 74 | 57.015 | 1.105 |  | 56.7 | 57.33 | 74 | 37274073
75 | 26.1 | 25.19 | -0.91 | -3.4866 | -3.5484 | 0 | 7.495 | 3.6024 | 75 | 25.645 | -3.5484 |  | 26.1 | 25.19 | 75 | 38170698
76 | 5.21 | 5.53 | 0.32 | 6.142 | 5.959 | 0 | 7.675 | -6.645 | 76 | 5.37 | 5.959 |  | 5.21 | 5.53 | 76 | 38171774
77 | 19.8 | 18.85 | -0.95 | -4.798 | -4.9159 | 0 | 7.755 | 4.2553 | 77 | 19.325 | -4.9159 |  | 19.8 | 18.85 | 77 | 38173004
78 | 1.35 | 1.38 | 0.03 | 2.2222 | 2.1978 | 0 | 7.775 | -8.1029 | 78 | 1.365 | 2.1978 |  | 1.35 | 1.38 | 78 | 32144703
79 | 8.15 | 8.76 | 0.61 | 7.4847 | 7.2147 | 0 | 7.96 | -7.2864 | 79 | 8.455 | 7.2147 |  | 8.15 | 8.76 | 79 | 32160631
80 | 7.36 | 7.63 | 0.27 | 3.6685 | 3.6024 | 0 | 8.205 | -4.2657 | 80 | 7.495 | 3.6024 |  | 7.36 | 7.63 | 80 | 32160629
81 | 81.4 | 75.25 | -6.15 | -7.5553 | -7.8519 | 0 | 8.455 | 7.2147 | 81 | 78.325 | -7.8519 |  | 81.4 | 75.25 | 81 | 32160628
82 | 31.8 | 33.75 | 1.95 | 6.1321 | 5.9497 | 0 | 8.475 | -5.7817 | 82 | 32.775 | 5.9497 |  | 31.8 | 33.75 | 82 | 38173363
83 | 21.2 | 18.5 | -2.7 | -12.7358 | -13.602 | 0 | 16.635 | -6.7929 | 83 | 19.85 | -13.602 |  | 21.2 | 18.5 | 83 | 32160632
84 | 24.2 | 19.72 | -4.48 | -18.5124 | -20.4007 | 0 | 16.77 | -14.6691 | 84 | 21.96 | -20.4007 |  | 24.2 | 19.72 | 84 | 38176246
85 | 5.23 | 4.7 | -0.53 | -10.1338 | -10.6747 | 0 | 19.325 | -4.9159 | 85 | 4.965 | -10.6747 |  | 5.23 | 4.7 | 85 | 32160633
86 | 6.75 | 6.69 | -0.06 | -0.88889 | -0.89286 | 0 | 19.85 | -13.602 | 86 | 6.72 | -0.89286 |  | 6.75 | 6.69 | 86 | 37284567
88 | 65 | 57.9 | -7.1 | -10.9231 | -11.5541 | 0 | 19.88 | 7.8471 | 87 | 61.45 | -11.5541 |  | 65 | 57.9 | 88 | 38175529
89 | 1.37 | 1.45 | 0.08 | 5.8394 | 5.6738 | 0 | 21.96 | -20.4007 | 88 | 1.41 | 5.6738 |  | 1.37 | 1.45 | 89 | 32003034
90 | 7.08 | 7.25 | 0.17 | 2.4011 | 2.3726 | 0 | 22.14 | -5.9621 | 89 | 7.165 | 2.3726 |  | 7.08 | 7.25 | 90 | 37287086
91 | 7.93 | 7.42 | -0.51 | -6.4313 | -6.645 | 0 | 25.645 | -3.5484 | 90 | 7.675 | -6.645 |  | 7.93 | 7.42 | 91 | 38176759
92 | 2.73 | 2.73 | 0 | 0 | 0 | 0 | 26.45 | -7.1834 | 91 | 2.73 | 0 |  | 2.73 | 2.73 | 92
93 | 2.04 | 1.98 | -0.06 | -2.9412 | -2.9851 | 0 | 32.775 | 5.9497 | 92 | 2.01 | -2.9851 |  | 2.04 | 1.98 | 93
94 | 4.44 | 4.36 | -0.08 | -1.8018 | -1.8182 | 0 | 36.74 | -23.19 | 93 | 4.4 | -1.8182 |  | 4.44 | 4.36 | 94
95 | 6.24 | 5.63 | -0.61 | -9.7756 | -10.278 | 0 | 39.61 | -2.979 | 94 | 5.935 | -10.278 |  | 6.24 | 5.63 | 95
96 | 3.83 | 3.58 | -0.25 | -6.5274 | -6.7476 | 0 | 45.03 | -7.8614 | 95 | 3.705 | -6.7476 |  | 3.83 | 3.58 | 96
97 | 3.01 | 2.43 | -0.58 | -19.2691 | -21.3235 | 0 | 47.215 | -11.3735 | 96 | 2.72 | -21.3235 |  | 3.01 | 2.43 | 97
98 | 2.36 | 2.2 | -0.16 | -6.7797 | -7.0175 | 0 | 47.31 | -11.7945 | 97 | 2.28 | -7.0175 |  | 2.36 | 2.2 | 98
99 | 3.61 | 3.07 | -0.54 | -14.9584 | -16.1677 | 0 | 54.39 | 1.8018 | 98 | 3.34 | -16.1677 |  | 3.61 | 3.07 | 99
100 | 6.11 | 6.58 | 0.47 | 7.6923 | 7.4074 | 0 | 57.015 | 1.105 | 99 | 6.345 | 7.4074 |  | 6.11 | 6.58 | 100
101 | 6.03 | 6.22 | 0.19 | 3.1509 | 3.102 | 0 | 58.2 | -26.4605 | 100 | 6.125 | 3.102 |  | 6.03 | 6.22 | 101
102 | 5.45 | 5.51 | 0.06 | 1.1009 | 1.0949 | 0 | 61.45 | -11.5541 | 101 | 5.48 | 1.0949 |  | 5.45 | 5.51 | 102
103 | 2.58 | 2.69 | 0.11 | 4.2636 | 4.1746 | 0 | 78.325 | -7.8519 | 102 | 2.635 | 4.1746 |  | 2.58 | 2.69 | 103
Note on A114: ###res
Data Assessment
P/B Regression
Y = 0.941 * X + 0.014
md(95) = 2.312
N = 102, r = 0.995
Cusum test for linearity shows no significant deviation from linearity.
* indicates significant difference (rejection of null-hypothesis,
gamma = 0.05 for slope or intercept from P/B and for slope from SPC).

## Result (hidden)
Method X: |  |  | Instrument:
Method Y: |  |  | Instrument:
0
Descriptive Statistics
| X | Y | Y-X | (Y - X)% of X
Median | 0 | 0 | 0 | 0
Mean | 0 | 0 | 0 | NV
Minimum | 0 | 0 | 0 | 0
Maximum | 0 | 0 | 0 | 0
68% Median Distance | 0 | 0 | 0 | 0
Standard Deviation | 0 | 0 | 0 | NV
Differences
Medians |  |  |  | 0
Means |  |  |  | 0
Regression and Correlation Analysis
Coefficients of Correlation: | 0 | 0
|  | slope b | intercept a | lower limit | upper limit
Structural Relationship Model:
Passing/Bablok (P/B) |  | 0 | 0
95% Confidence Region for b (P/B) |  |  |  | 0 | 0
95% Confidence Region for a (P/B) |  |  |  | 0 | 0
Std. Principal Component (SPC) |  | 0 | 0
Linear Model:
Least Squares Regression |  | 0 | 0
Theil Regression |  | 0 | 0
Dispersion of Residuals:
Passing/Bablok Regression | md(68) = | 0 | md(95) = | 0
Std. Principal Component |  |  | SE = | 0
Statistics | Method X: | Method Y:
N | 0 | 0
Mean | 0 | 0
Median | 0 | 0
Minimum | 0 | 0
Maximum | 0 | 0
Range | 0 | 0
Difference Plot
(Normalized Differences)
| Sample Values | [18 empty] | 5 | 10 | [5 empty] | 5 | 20
Serial Number | X | Y | Y - X | (Y - X)% of X | Normalized Difference (%) | [14 empty] | 10 | 15 | [5 empty] | 10 | 25
[6 empty] | Diff Data Preset |  | Diff Ref Preset
scale_min | 0 | 0 | 0 |  |  | 0.4 | 0.05 | 0 | -0.05 |  |  |  | SelectData | [6 empty] | SelectCalibs | [6 empty] | SelectQCs | [6 empty] | F1_select |  | F2_select |  | grp_select
scale_max | 100 |  | 80 |  |  | 0.6 | -0.05 | 1 | -0.05 |  |  |  | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? | [6 empty] | #NAME? |  | #NAME?
| 0 |  |  | 5 | [7 empty] | data_count | [6 empty] | calib_count | [6 empty] | QC_count | [6 empty] | F1_count |  | F2_count |  | selFilt_count |  |  |  |  | selCrit |  |  | H | I | J | K | L | M
| 100 |  |  | 95 | [8 empty] | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2 |  | filter1 |  | filter2 |  | Selected Filters | F1 | Grp Start | Legend |  | X_Vektor | Y_Vektor | X_Vektor | Y_Vektor | lfd_no | ident_no | filter1 | filter2
1 | 24.5 | 26.2 | 1.7 | 6.9388 | 6.7061 | 0 | 0 | 0 | 1 | 25.35 | 6.7061 | [8 empty] | -5 | 10 | [5 empty] | -5 | 20
2 | [8 empty] | 2 | [10 empty] | -10 | 15 | [5 empty] | -10 | 25
Note on A114: ###res
Data Assessment